System Requirements for Non-Contact Type Clinical Thermometer for Adult People

Part Number: 24-06-TRM-801

Rev: 02

Date: 8 Oct 2024

ALL RIGHTS RESERVED. REPRODUCTION OR ISSUE TO THIRD PARTIES IN ANY FORM WHATEVER IS NOT PERMITTED WITHOUT WRITTEN AUTHORITY FROM THE PROPRIETORS.

# CHANGE RECORDS

FIRST RELEASE

CREATED BY

Name	Title	Date	Sign.

Beyza YAŞAR	Electric Eng                                                                             08.10.2024

CHECKED BY

#### Name	Title	Date	Sign.

Doğa ERPAK	Quality Eng                                                                            08.10.2024 08.10.2024

APPROVED BY

#### Name	Title	Date	Sign.

Elif AĞILLAR                   Program Mng	08.10.2024

### REVISIONS

Rev. No	Change Description	Date	QA Approval

01                 Purpose, Units of Measure, Measuring Site                       04.11.2024          Approved

02                Measuring Site, Operating Controls                                            19.11.2024              Approved

# TABLE OF CONTENTS

CHANGE RECORDS	2

FIRST RELEASE	2

REVISIONS	2

- PURPOSE	4

- SCOPE	4

- REFERENCES	4

- REQUIREMENTS	5

  - PHYSICAL REQUİREMENTS	5

    - WEIGHT	5

    - SURFACE	5

  - FUNCTIONAL REQUİREMENTS	5

    - OPERATING CONTROLS	5

    - POWER SUPPLY	5

    - UNITS OF MEASURE	5

    - RATED OUTPUT RANGE	5

    - ACCURACY	6

    - MEASUREMENT TIME	6

    - TECHNICAL ALARM CONDITIONS	6

  - ENVIRONENTAL CONDITIONS	6

    - OPERATING TEMPERATURE	6

    - STORAGE TEMPERATURE	6

  - USABILITY	6

    - DISPLAY SIZE	6

    - MEASURING SITE	7

## PURPOSE

This document outlines the functional, performance, safety, quality, maintainability, and reliability requirements for the ongoing development of the Non-Contact Type Clinical Thermometer for adult people.

## SCOPE

The identification, definition, and documentation of basic safety and essential performance requirements are crucial for the development of a Non-Contact type clinical thermometer tailored for elderly individuals. This encompasses the selection of technology, both fundamental and optional functionalities, environmental considerations, and performance criteria. The specified requirements may encompass built-in tests, testability, external interfaces, environmental factors, testing and maintenance considerations, power requirements, and physical characteristics.

Utilization of requirements standards during the requirements capture phase aids in establishing rules, procedures, methods, guidance, and criteria for developing these specifications. These standards also encompass methodologies and qualification tests designed to facilitate the verification, validation, and certification processes.

## REFERENCES

  - IEC 60601–1:2005+A1:2012 : Medical electrical equipment - Part 1: General requirements for basic safety and essential performance

  - IEC 60601‐1‐2:2014 : Medical electrical equipment — Part 1-2: General requirements for basic safety and essential performance — Collateral Standard: Electromagnetic disturbances Requirements and tests

  - ISO 80601-2-56:2017 : Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement

## REQUIREMENTS

## PHYSICAL REQUİREMENTS

#### WEIGHT

The weight of the thermometer shall be less than 200 grams.

#### SURFACE

The clinical thermometer and its sensor shall be designed to allow for surface cleaning and disinfection to reduce to acceptable levels the risk of infection of operators, bystanders, or patients.

## FUNCTIONAL REQUİREMENTS

#### OPERATING CONTROLS

There shall be a momentary switch (ON Switch) to start the measurement and automatically be turned off after 30 seconds when the measurement is completed.

There shall be a MODE switch to select the mode of operation and have a memory of at least 3 data. It will also be designed to switch between Fahrenheit and Celsius units. (adjusted, test, etc.). There should be an indicator on the thermometer's digital display that shows the battery charge status.

#### POWER SUPPLY

The equipment shall have a rechargeable power unit which can be charged by 5V cell phone adaptors through USB-C.

#### UNITS OF MEASURE

For the Non-Contact Type clinical thermometer for elderly people, it is decided to express temperature in degrees Celsius, ℃. Additionally, the main unit of measurement, Celsius, can be converted to Fahrenheit.

#### RATED OUTPUT RANGE

The output temperature of the clinical thermometer should cover the minimum rated output range from 34,0 °C to 43,0 °C. Compliance is checked by inspection and functional testing.

#### ACCURACY

The LABORATORY ACCURACY within the RATED OUTPUT RANGE in NORMAL USE shall be within

±0,3°C.

#### MEASUREMENT TIME

The measurement time of a non-contact thermometer must range between 1 to 3 seconds. The exact time depends on the sensor's performance, processing speed, and environmental conditions.

#### ALARM CONDITIONS

The thermometer shall provide an ALARM sound and/or display for the following:

      - Completion of measurement

      - Low voltage of the internal electrical power source

      - Output temperature outside the rated output range.

## ENVIRONMENTAL CONDITIONS

#### OPERATING TEMPERATURE

The device shall meet the stated accuracy requirement when operating +15 to +40℃ temperature range.

#### STORAGE TEMPERATURE

The device shall meet the stated accuracy requirement after storing for the duration of 4 hours at

-20 to +60℃ temperature range.

## USABILITY

#### DISPLAY SIZE

#### The display of the output thermometer shall be at least 4 mm high or optically magnified to appear at that height.

#### MEASURING SITE

The Non-Contact thermometer must maintain the specified distance from the measuring site and this distance to the measuring site for a  Non-Contact thermometer ranges between 3 to 5 cm.