BS EN ISO 80601-2-56:2017

BSI Standards Publication

Medical electrical equipment

Part 2-56: Particular requirements for basic safety and essential performance of  clinical thermometers for body temperature measurement (ISO 80601-2-56:2017)

BS EN ISO 80601-2-56:2017 BRITISH STANDARD

National foreword

This British Standard is the UK implementation of

EN ISO 80601‑2‑56:2017. It is identical to ISO 80601‑2‑56:2017.

It supersedes BS EN ISO 80601‑2‑56:2012, which is withdrawn.

The UK participation in its preparation was entrusted to Technical

Committee CH/121/9, Lung Ventilators & Related Equipment.

A list of organizations represented on this committee can be obtained on

request to its secretary.

This publication does not purport to include all the necessary provisions

of a contract. Users are responsible for its correct application.

© The British Standards Institution 2017

Published by BSI Standards Limited 2017

ISBN 978 0 580 87921 0

ICS 11.040.55

Compliance with a British Standard cannot confer immunity from

legal obligations.

This British Standard was published under the authority of the

Standards Policy and Strategy Committee on 30 September 2017.

Amendments/corrigenda issued since publication

Date Text affected

EUROPEAN STANDARD  NORME EUROPÉENNE  EUROPÄISCHE NORM

BS EN ISO 80601-2-56:2017

EN ISO 80601-2-56   July 2017

ICS 11.040.55 Supersedes EN ISO 80601-2-56:2012

English Version

Medical electrical equipment - Part 2-56: Particular

requirements for basic safety and essential performance of clinical thermometers for body temperature measurement  (ISO 80601-2-56:2017)

Appareils électromédicaux - Partie 2-56: Exigences  particulières relatives à la sécurité fondamentale et aux  performances essentielles des thermomètres médicaux

pour mesurer la température de corps (ISO 80601-2- 56:2017)

This European Standard was approved by CEN on 28 June 2017.

Medizinische elektrische Geräte - Teil 2-56: Besondere  Festlegungen für die Sicherheit einschließlich der  wesentlichen Leistungsmerkmale von medizinischen  Thermometern zum Messen der Körpertemperatur  (ISO 80601-2-56:2017)

CEN members are bound to comply with the CEN/CENELEC Internal Regulations which stipulate the conditions for giving this  European Standard the status of a national standard without any alteration. Up-to-date lists and bibliographical references  concerning such national standards may be obtained on application to the CEN-CENELEC Management Centre or to any CEN  member.

This European Standard exists in three official versions (English, French, German). A version in any other language made by  translation under the responsibility of a CEN member into its own language and notified to the CEN-CENELEC Management  Centre has the same status as the official versions.

CEN members are the national standards bodies of Austria, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Estonia,  Finland, Former Yugoslav Republic of Macedonia, France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Latvia, Lithuania,  Luxembourg, Malta, Netherlands, Norway, Poland, Portugal, Romania, Serbia, Slovakia, Slovenia, Spain, Sweden, Switzerland,  Turkey and United Kingdom.

EUROPEAN COMMITTEE FOR STANDARDIZATION

C O M I T É E U R O P É E N D E N O R M A L I S A T I O N

E U R O P Ä I S C H E S K O M I T E E F Ü R N O R M U N G

CEN-CENELEC Management Centre: Avenue Marnix 17, B-1000 Brussels

© 2017 CEN All rights of exploitation in any form and by any means reserved  worldwide for CEN national Members.

Ref. No. EN ISO 80601-2-56:2017 E

BS EN ISO 80601-2-56:2017European foreword

EN ISO 80601-2-56:2017 (E)

This document (EN ISO 80601-2-56:2017) has been prepared by Technical Committee  ISO/TC 121 “Anaesthetic and respiratory equipment” in collaboration with Technical Committee  CEN/TC 205 “Non-active medical devices” the secretariat of which is held by DIN.

This European Standard shall be given the status of a national standard, either by publication of an  identical text or by endorsement, at the latest by January 2018, and conflicting national standards shall  be withdrawn at the latest by July 2020.

Attention is drawn to the possibility that some of the elements of this document may be the subject of  patent rights. CEN shall not be held responsible for identifying any or all such patent rights.

This document supersedes EN ISO 80601-2-56:2012.

This document has been prepared under a mandate given to CEN by the European Commission and the  European Free Trade Association, and supports essential requirements of EU Directive(s).

For relationship with EU Directive(s), see informative Annex ZA, which is an integral part of this  document.

According to the CEN-CENELEC Internal Regulations, the national standards organizations of the  following countries are bound to implement this European Standard: Austria, Belgium, Bulgaria,  Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, Former Yugoslav Republic of Macedonia,  France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Latvia, Lithuania, Luxembourg, Malta,  Netherlands, Norway, Poland, Portugal, Romania, Serbia, Slovakia, Slovenia, Spain, Sweden, Switzerland,  Turkey and the United Kingdom.

The following referenced documents are indispensable for the application of this document. For  undated references, the latest edition of the referenced document (including any amendments) applies.  For dated references, only the edition cited applies. However, for any use of this standard ‘within the  meaning of Annex ZA’, the user should always check that any referenced document has not been  superseded and that its relevant contents can still be considered the generally acknowledged state-of

art.

When an IEC or ISO standard is referred to in the ISO standard text, this shall be understood as a  normative reference to the corresponding EN standard, if available, and otherwise to the dated version  of the ISO or IEC standard, as listed below.

NOTE The way in which these referenced documents are cited in normative requirements determines the  extent (in whole or in part) to which they apply.

3

EN ISO 80601-2-56:2017 (E)

BS EN ISO 80601-2-56:2017

Table 1 — Correlations between undated normative references and dated EN and ISO standards

| Normative references as listed in Clause 201.2 of the ISO standard | Equivalent dated standard EN ISO or IEC |  |
| --- | --- | --- |
| IEC 60601-1 | EN 60601-1:2006 + Cor. :2010 + A1:2013 | IEC 60601-1:2005 + Cor. :2006 + Cor. :2007 + A1:2012 |
| IEC 60601-1-2 | EN 60601-1-2:2015 | IEC 60601-1-2:2014 |
| IEC 60601-1-6 | EN 60601-1-6:2010 + A1:2015 | IEC 60601-1-6:2010 + A1:2013 |
| IEC 60601-1-8 | EN 60601-1-8:2007 + Cor.:2010 + A1:2013 | IEC 60601-1-8:2006 + A1:2012 |
| IEC 60601-1-11 | EN 60601-1- 11:2015 | IEC 60601-1-11:2015 |
| IEC 60601-1-12 | EN 60601-1- 12:2015 | IEC 60601-1-12:2014 |
| IEC 62366-1 | EN 62366-1:2015 | IEC 62366-1:2015 |
| ISO 14155:2011 | EN ISO 14155:2011 + AC:2011 | ISO 14155:2011 + AC:2011 |
| ISO 14937 | EN ISO 14937:2009 | ISO 14937:2009 |
| ISO 15223-1 | EN ISO 15223- 1:2016 | ISO 15223-1:2016 |
| ISO 17664 | EN ISO 17664:2004 | ISO 17664:2004 |

Endorsement notice

The text of ISO 80601-2-56:2017 has been approved by CEN as EN ISO 80601-2-56:2017 without any  modification.

4

BS EN ISO 80601-2-56:2017

Annex ZA (informative)

EN ISO 80601-2-56:2017 (E)

Relationship between this European Standard and the essential  requirements of Directive 93/42/EEC [OJ L 169] aimed to be covered

This European Standard has been prepared under a Commission’s standardization request M/023  concerning the development of European Standards related to medical devices to provide one voluntary  means of conforming to essential requirements of Council Directive 93/42/EEC of 14 June 1993  concerning medical devices [OJ L 169].

Once this standard is cited in the Official Journal of the European Union under that Directive,  compliance with the normative clauses of this standard given in Table ZA.1 confers, within the limits of  the scope of this standard, a presumption of conformity with the corresponding essential requirements  of that Directive and associated EFTA regulations.

NOTE 1 Where a reference from a clause of this standard to the risk management process is made, the risk  management process needs to be in compliance with Directive 93/42/EEC as amended by 2007/47/EC. This  means that risks have to be reduced ‘as far as possible’, ‘to a minimum’, ‘to the lowest possible level’, ‘minimized’  or ‘removed’, according to the wording of the corresponding essential requirement.

NOTE 2 The manufacturer’s policy for determining acceptable risk must be in compliance with Essential  Requirements 1, 2, 5, 6, 7, 8, 9, 11 and 12 of the Directive.

NOTE 3 This Annex ZA is based on normative references according to the table of references in the European  foreword, replacing the references in the core text.

NOTE 4 When an Essential Requirement does not appear in Table ZA.1, it means that it is not addressed by this  European Standard.

Table ZA.1 — Correspondence between this European Standard and Annex I of Directive  93/42/EEC [OJ L 169]

| Essential Requirements of Directive 93/42/EEC | Clause(s)/sub-clause(s) of this EN | Remarks/Notes |
| --- | --- | --- |
| 7.2 | 201.11.6.6 | ER 7.2 is covered only in respect to the patient, operator and user. |
| 8.7 | 201.7.2.1.101 c) |  |
| 9.1 | 201.7.9.2.101 f), 201.16, 201.101.1 2nd para, 201.102.1 3rd para, 201.103, 201.103.2 | ER 9.1 is covered by 201.103 in respect of probes, probe cable extenders and probe covers only. |
| 9.2 | 201.9, 201.12.1.101, 201.12.2, 201.15, 202 | ER 9.2 is covered by the listed standard clauses as follows: — 201.9 to the extent set out in the specified EN version of IEC 60601-1:2005+A1, Clause 9, — 201.12.1.101 for accuracy of controls and |

5

EN ISO 80601-2-56:2017 (E)

BS EN ISO 80601-2-56:2017

| Essential Requirements of Directive 93/42/EEC | Clause(s)/sub-clause(s) of this EN | Remarks/Notes |
| --- | --- | --- |
|  |  | instruments only, — 201.12.2 for display size and display integrity only, — 201.15 to the extent set out in the specified EN version of IEC 60601-1:2005+A1, Clause 15, — 202 Electromagnetic Compatibility. The 4th indent of ER 9.2 is not covered. |
| 9.3 | 201.13 | ER 9.3 is covered by 201.13 to the extent set out in the specified EN version of IEC 60601- 1:2005+A1:2012, Clause 13 only. |
| 10.1 | 201.7.9.2.101 d), 201.7.9.2.101 e), 201.12, 201.101, 201.102, 201.103 |  |
| 10.2 | 201.12.2 |  |
| 10.3 | 201.7 | ER 10.3 is covered by 201.7.4.3.101 and 201.7.9.2.101 i) |
| 11.3.1 | 202 |  |
| 12.1 | 201.14 | ER 12.1 is covered by 201.14 to the extent set out in the specified EN version of IEC 60601- 1:2005 +A1:2012, Clause 14 only. |
| 12.1 a) | 201.14 | ER 12.1 a) is covered by 201.14 to the extent set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 14 only. |
| 12.4 | 201.12 | ER 12.4 is covered by 201.12 to the extent set out in the specified EN version of IEC 60601- 1:2005 +A1:2012, Clause 12 only. Directive Annex 1 ER 12.4 is only covered if the alarm alerts the user to situations which could lead to death or severe deterioration of the patient's state of health. |
| 12.5 | 202 |  |
| 12.6 | 201.8 | ER 12.6 is covered by 201.8 to the extent set out in the specified EN version of IEC 60601- 1:2005 +A1:2012, Clause 8 only. |
| 12.7.1 | 201.9 | ER 12.7.1 is covered by 201.9 to the extent set out in the specified EN version of IEC 60601- 1:2005 +A1:2012, Clause 9 only. |
| 12.7.2 | 201.9 | ER 12.7.2 is covered by 201.9 to the extent set out in the specified EN version of IEC 60601- 1:2005 +A1:2012, Clause 9 only. |

6

BS EN ISO 80601-2-56:2017

EN ISO 80601-2-56:2017 (E)

| Essential Requirements of Directive 93/42/EEC | Clause(s)/sub-clause(s) of this EN | Remarks/Notes |
| --- | --- | --- |
| 12.7.3 | 201.9 | ER 12.7.3 is covered by 201.9 to the extent set out in the specified EN version of IEC 60601- 1:2005 +A1:2012, Clause 9 only. |
| 12.7.4 | 201.8, 201.11, 201.15 | ER 12.7.4 is covered — by 201.8 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 8 only, — by 201.11 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 11 only, — by 201.15 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 15 only. |
| 12.7.5 | 201.11, 201.15 | ER 12.7.4 is covered — by 201.11 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 11 only, — by 201.15 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 15 only. |
| 12.9 | 201.7, 201.12.2, 201.15, 206 |  |
| 13.1 | 201.7, 201.7.2.1, 201.7.2.1.101, 201.7.2.2, 201.7.9 | The requirement for information on the sales packaging is not addressed. The last para of ER 13.1 is not covered. |
| 13.2 | 201.7, 201.7.2.1, 201.8, 201.9 | ER 13.2 is covered provided that any used symbols conform to harmonized standards and where no harmonized standards exist, the symbols and colours are specified in the instructions for use. |
| 13.3 b) | 201.7, 201.7.2.1.101 b) | ER 13.3 b) is covered by 201.7 as set out in the specified EN version of IEC 60601- 1:2005+A1:2012, Clause 7.2.2 only. |
| 13.3 c) | 201.7.2.1.101 c) | ER 13.3 c) is covered by 201.7.2.1.101 provided that the information specified appears on the device label. |
| 13.3 e) | 201.7.2.1.101 d) | ER 13.3 e) is covered by 201.7.2.1.101 provided that the information appears on the device label and that the 'Use by' date is expressed as a year and month in that order. |
| 13.3 i) | 201.7, 201.7.2.1.101 e) | ER 13.3 i) is covered by 201.7 as set out in the specified EN version of IEC 60601- 1:2005+A1:2012, Clause 7 only. |

7

EN ISO 80601-2-56:2017 (E)

BS EN ISO 80601-2-56:2017

| Essential Requirements of Directive 93/42/EEC | Clause(s)/sub-clause(s) of this EN | Remarks/Notes |
| --- | --- | --- |
|  |  | ER 13.3 i) is covered by 201.7.2.1.101 e) provided that the information appears on the device label. |
| 13.3 j) | 201.7 | ER 13.3 j) is covered by 201.7 to the extent set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 7 only and provided that the information is shown on the device label. |
| 13.3 k) | 201.7 | ER 13.3 k) is covered by 201.7 to the extent set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 7 only and provided that the information is shown on the device label. |
| 13.3 m) | 201.7.2.1.101 c) | Presumption of conformity is only provided if symbols 4 to 7 are utilized. ER 13.3 m) is covered by 201.7.2.1.101 c) provided that the information is shown on the device label. |
| 13.6 a) | 201.7, 201.7.9.1, 201.16 | ER 13.6 a) is covered by 201.7, 201.7.9.1, and 201.16 provided that all relevant information in ER 13.3 a) to c) and f) to n) are covered in the instructions for use. |
| 13.6 b) | 201.7.9.2.101 d), 201.7.9.2.101 e) | ER 13.6 b) is covered by 201.7.9.2.101 d) and 201.7.9.2.101 e) for the parameters listed in these clauses only. |
| 13.6 c) | 201.7, 201.7.9.2.101 f), 201.16 | ER 13.6 c) is covered by — 201.7 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 7 only, provided the information is included in the instructions for use, — 201.7.9.2.101 f) for the information set out in this clause only, — 201.16 as set out in the specified EN version of IEC 60601-1:2005 +A1:2012, Clause 16 only for the parameters listed in that standard provided the information is included in the instructions for use. |
| 13.6 d) | 201.7, 201.7.9.2.101 h), 201.16 | ER 13.6 d) is covered by — 201.7 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 7 only, provided the information is included in the instructions for use, — 201.7.9.2.101 h) for the information set out in this clause only, |

8

BS EN ISO 80601-2-56:2017

EN ISO 80601-2-56:2017 (E)

| Essential Requirements of Directive 93/42/EEC | Clause(s)/sub-clause(s) of this EN | Remarks/Notes |
| --- | --- | --- |
|  |  | — 201.16 as set out in the specified EN version of IEC 60601-1:2005 +A1:2012, Clause 16 only for the parameters listed in that standard provided the information is included in the instructions for use. |
| 13.6 f) | 201.7, 201.16 | ER 13.6 f) is covered by — 201.7 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 7 only, provided the information is included in the instructions for use, — 201.16 as set out in the specified EN version of IEC 60601-1:2005 +A1:2012, Clause 16 only for the parameters listed in that standard provided the information is included in the instructions for use. |
| 13.6 h) | 201.7, 201.7.2.9.2.101 k), 201.11, 201.16 | The requirement for information on the packaging is not addressed. ER 13.6 h) is covered by — 201.7 as set out in the specified EN version of IEC 60601-1:2005+A1:2012, Clause 7 only, provided the information is included in the instructions for use, — 201.7.9.2.101 k) for the information set out in this clause only, — 201.16 as set out in the specified EN version of IEC 60601-1:2005 +A1:2012, Clause 16 only for the parameters listed in that standard provided the information is included in the instructions for use. |
| 13.6 i) | 201.7 | ER 13.6 i) is covered by 201.7 as set out in the specified EN version of IEC 60601- 1:2005+A1:2012, Clause 7 only, provided the information is included in the instructions for use. |
| 13.6 n) | 201.7.9.2.101 j) |  |
| 13.6 p) | 201.7.9.2.101 e) |  |

WARNING 1 — Presumption of conformity stays valid only as long as a reference to this European  Standard is maintained in the list published in the Official Journal of the European Union. Users of this  standard should consult frequently the latest list published in the Official Journal of the European  Union.

WARNING 2 — Other Union legislation may be applicable to the products falling within the scope of  this standard.

9

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Contents Page

Foreword ........................................................................................................................................................................... v Introduction ................................................................................................................................................................... vi 201.1 * Scope, object and related standards ...................................................................................................... 1 201.1.1 Scope ............................................................................................................................................................... 1 201.1.2 Object .............................................................................................................................................................. 1 201.1.3 Collateral standards .................................................................................................................................. 2 201.1.4 Particular standards ................................................................................................................................. 2 201.2 Normative references .................................................................................................................................... 3 201.3 Terms and definitions .................................................................................................................................... 4 201.4 General requirements .................................................................................................................................... 7 201.4.2 RISK MANAGEMENT PROCESS for ME EQUIPMENT or ME SYSTEMS ............................................................ 7 201.4.3 ESSENTIAL PERFORMANCE .............................................................................................................................. 8 Table 201.101 — Distributed ESSENTIAL PERFORMANCE requirements .......................................................... 8 201.5 General requirements for testing of ME EQUIPMENT .............................................................................. 8 201.6 Classification of ME EQUIPMENT and ME SYSTEMS ....................................................................................... 8 201.7 ME EQUIPMENT identification, marking and documents ...................................................................... 8 201.7.9 ACCOMPANYING DOCUMENT ........................................................................................................................... 9 201.8 Protection against electrical HAZARDS from ME EQUIPMENT .............................................................. 10 201.9 Protection against mechanical HAZARDS of ME EQUIPMENT and ME SYSTEMS ................................. 10 201.10Protection against unwanted and excessive radiation HAZARDS .................................................. 10 201.11Protection against excessive temperatures and other HAZARDS .................................................. 11 201.11.7 Biocompatibility of ME EQUIPMENT and ME SYSTEMS ..................................................................... 11 201.12Accuracy of controls and instruments and protection against hazardous outputs .............. 12 201.12.1 Accuracy of controls and instruments ..........................................................................................12 201.12.2 Usability ................................................................................................................................................... 12 201.13HAZARDOUS SITUATIONS and fault conditions.......................................................................................... 12 201.14PROGRAMMABLE ELECTRICAL MEDICAL SYSTEMS (PEMS) ............................................................................. 12 201.15Construction of ME EQUIPMENT ................................................................................................................... 12 201.16ME SYSTEMS ....................................................................................................................................................... 13 201.17Electromagnetic compatibility of ME EQUIPMENT and ME SYSTEMS .................................................. 13 201.101 Laboratory performance requirements ....................................................................................... 13 201.101.1 * General test requirements ....................................................................................................... 13 201.101.2 * LABORATORY ACCURACY .................................................................................................................. 13

© ISO 2017 – All rights reserved iii

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.101.3 * Time response for a DIRECT MODE CLINICAL THERMOMETER that is not an ADJUSTED MODE CLINICAL THERMOMETER ..................................................................................................... 14

201.102 * CLINICAL ACCURACY VALIDATION ......................................................................................................... 15 201.102.1 Method ............................................................................................................................................... 15 201.102.2 * Human subject population requirements .......................................................................... 16 Table 201.102 — Subject age groups .................................................................................................................. 16 201.102.3 * CLINICAL BIAS calculation ............................................................................................................ 17 201.102.4 * LIMITS OF AGREEMENT calculation ............................................................................................. 17 201.102.5 * CLINICAL REPEATABILITY calculation ......................................................................................... 18 201.103 * PROBES, PROBE CABLE EXTENDERS and PROBE COVERS ................................................................... 18 201.103.1 General ............................................................................................................................................... 18 201.103.2 Labelling ............................................................................................................................................ 19 202 Electromagnetic disturbances — Requirements and tests ........................................................... 19 206 Usability ........................................................................................................................................................... 20

208 General requirements, tests and guidance for alarm systems in medical electrical  equipment and medical electrical systems......................................................................................... 20

211 Requirements for medical electrical equipment and medical electrical systems used  in the home healthcare environment ................................................................................................... 21

212 Requirements for medical electrical equipment and medical electrical systems  intended for use in the emergency medical services environment ........................................... 21

Annex C (informative) Guide to marking and labelling requirements for ME EQUIPMENT and  ME SYSTEMS ........................................................................................................................................................ 22

Annex D (informative) Symbols on marking................................................................................................. 25 Annex AA (informative) Particular guidance and rationale.................................................................... 27

Figure AA.101 — Example of temperature time adjustment for a predictive intermittent  CLINICAL THERMOMETER .................................................................................................................................. 29

Figure AA.102 — General structure of a CLINICAL THERMOMETER ................................................................. 30 Table AA.101 — Required tests for CLINICAL THERMOMETERS ........................................................................ 33

Table AA.102 — Example combinations of operating conditions and REFERENCE temperature  for testing the LABORATORY ACCURACY ...................................................................................................... 34

Figure AA.103 — Example of a comparison plot for DUT and RCT ............................................................ 36

Figure AA.104 — Example of a Bland-Altman Plot[18] of the temperature difference (DUT  minus RCT) versus the average OUTPUT TEMPERATURES of two thermometers .......................... 37

Table AA.103 — Example of CLINICAL ACCURACY VALIDATION test results .................................................... 38 Annex BB (informative) Reference temperature source.......................................................................... 40

Annex CC (informative) Reference to the essential principles of safety and performance of  medical devices in accordance with ISO 16142-1[24] ....................................................................... 42

Table CC.1 — Correspondence between the essential principles and this document ....................... 42 Annex DD (informative) Terminology — Alphabetized index of defined terms.............................. 45 Bibliography ................................................................................................................................................................. 48

iv © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Foreword

ISO  (the  International  Organization  for  Standardization)  is  a  worldwide  federation  of  national  standards  bodies  (ISO  member  bodies).  The  work  of  preparing  International  Standards  is  normally  carried  out  through  ISO  technical  committees. Each member  body interested in a  subject  for which a  technical  committee  has  been  established  has  the  right  to  be  represented  on  that  committee.  International organizations, governmental and non‐governmental, in liaison with ISO, also take part in  the  work.  ISO  collaborates  closely  with  the  International  Electrotechnical  Commission  (IEC)  on  all  matters of electrotechnical standardization.

The  procedures  used  to  develop  this  document  and  those  intended  for  its  further  maintenance  are  described in  the  ISO/IEC Directives, Part 1.  In particular  the different approval criteria needed  for  the  different  types of  ISO documents should be noted. This document was drafted in accordance with  the  editorial rules of the ISO/IEC Directives, Part 2 (see www.iso.org/directives).

Attention is drawn to the possibility that some of the elements of this document may be the subject of  patent rights.  ISO shall not be held responsible  for identifying any or all such patent rights. Details of  any patent rights identified during the development of the document will be in the Introduction and/or  on the ISO list of patent declarations received (see www.iso.org/patents).

Any trade name used in this document is information given for the convenience of users and does not  constitute an endorsement.

For  an  explanation  on  the  meaning  of  ISO  specific  terms  and  expressions  related  to  conformity  assessment,  as  well  as  information  about  ISO's  adherence  to  the  World  Trade  Organization  (WTO)  principles  in  the  Technical  Barriers  to  Trade  (TBT)  see  the  following  URL:  www.iso.org/iso/foreword.html.

The  committee  responsible for  this  document  is  ISO/TC 121,  Anaesthetic and respiratory equipment,  Subcommittee  SC 3,  Lung ventilators and related equipment,  and  Technical  Committee  IEC/TC 62,  Electrical equipment in medical practice, Subcommittee SC D, Electrical equipment.

This  second  edition  cancels  and  replaces  the  first  edition  (ISO 80601‐2‐56:2009),  which  has  been  technically  revised.  It  also  incorporates  the  Amendments  IEC 60601‐1:2005/AMD1:2012,  IEC 60601‐1‐6:2010/AMD1:2013 and IEC 60601‐1‐8:2006/AMD1:2012, as well as IEC 60601‐1‐12, the  second edition of IEC 60601‐1‐11 and the fourth edition of IEC 60601‐1‐2.

The most significant changes are the following modifications:

—  change in the clinical evaluation exclusion criteria related to antipyretics;

—  deletion of Annex CC as this material is covered by IEC 60601‐1‐9[1];

and the following additions:

—  disclosure requirement for a summary of the USE SPECIFICATION;

—  tests for mechanical strength (via IEC 60601‐1‐11 and IEC 60601‐1‐12);

—  tests for ENCLOSURE integrity (water ingress via IEC 60601‐1‐11 and IEC 60601‐1‐12);  —  tests for cleaning and disinfection PROCEDURES (via IEC 60601‐1‐11 and IEC 60601‐1‐12).

© ISO 2017 – All rights reserved v

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Introduction

This document deals with electrical CLINICAL  THERMOMETERS, either already available or  that will come  available in the future.

The purpose of a CLINICAL THERMOMETER is to assess the true temperature of a REFERENCE BODY SITE. The  temperature  of  the  PATIENT'S  body  is  an  important  vital  sign  in  assessing  overall  health,  typically  in  combination with blood pressure and pulse rate. Determining whether a PATIENT is afebrile,  febrile or  hypothermic is an important purpose of a CLINICAL  THERMOMETER, since being  febrile suggests  that  the  PATIENT is ill.

There  are  different  temperatures  at  each  REFERENCE  BODY  SITE  according  to  the  balance  between  the  production,  transfer,  and  loss  of  heat[2].  CLINICAL  ACCURACY  of  a  CLINICAL  THERMOMETER  is  VERIFIED  by  comparing  its  OUTPUT  TEMPERATURE  with  that  of  a  REFERENCE  THERMOMETER,  which  has  a  specified  uncertainty  for  measuring  true  temperature.  For  an  equilibrium  CLINICAL  THERMOMETER,  the  CLINICAL  ACCURACY  can  be  sufficiently  determined  under laboratory  conditions  that  create an equilibrium  state  between the two thermometers.

For  a  CLINICAL  THERMOMETER  that  operates  in  the  ADJUSTED  MODE,  laboratory  VERIFICATION  alone  is  not  sufficient  because  the  adjustment  algorithm  for  deriving  the  OUTPUT  TEMPERATURE  includes  the  characteristics  of  the  PATIENT  and  the  environment[3].  Therefore,  the  CLINICAL  ACCURACY  of  a  CLINICAL  THERMOMETER that operates in the ADJUSTED MODE has to be VALIDATED clinically, using statistical methods  of  comparing  its  OUTPUT  TEMPERATURE  with  that  of  a  REFERENCE  CLINICAL  THERMOMETER  which  has  a  specified CLINICAL ACCURACY in representing a particular REFERENCE BODY SITE temperature.

For a CLINICAL THERMOMETER that operates in the ADJUSTED MODE, the LABORATORY ACCURACY is VERIFIED in a  DIRECT  MODE  and  the  CLINICAL  ACCURACY  is  VALIDATED  in  the  ADJUSTED  MODE  (OPERATING  MODE)  with  a  sufficiently large group of human subjects.

The  intention  of  this  document  is  to  specify  the  requirements  and  the  test  PROCEDURES  for  the  VERIFICATION of the LABORATORY ACCURACY for all types of electrical CLINICAL THERMOMETERS as well as for  the VALIDATION of the CLINICAL ACCURACY of a CLINICAL THERMOMETER that operates in the ADJUSTED MODE.

This document has been drafted in accordance with the ISO/IEC Directives, Part 2.  In this document, the following print types are used.

—  Requirements and definitions: roman type.

—  Test specifications: italic type.

—  Informative material appearing outside of tables, such as notes, examples and references: in smaller  type. Normative text of tables is also in a smaller type.

—  TERMS DEFINED IN CLAUSE 3 OF THE GENERAL STANDARD, IN THIS DOCUMENT OR AS NOTED: SMALL CAPITALS.  In referring to the structure of this document, the term

—  “clause”  means  one  of  the  numbered  divisions  within  the  table  of  contents,  inclusive  of  all  subdivisions (e.g. Clause 7 includes subclauses 7.1, 7.2, etc.), and

—  “subclause” means a numbered subdivision of a clause (e.g. 7.1, 7.2 and 7.2.1 are all subclauses of  Clause 7).

vi © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

References  to clauses within  this document are preceded by  the term  “Clause”  followed by  the clause  number. References to subclauses within this document are by number only.

In  this  document,  the  conjunctive  “or”  is  used  as  an  “inclusive  or”  so  a  statement  is  true  if  any  combination of the conditions is true.

The  verbal  forms  used  in  this  document  conform  to  usage  described  in  Annex H  of  the  ISO/IEC  Directives, Part 2. For the purposes of this document, the auxiliary verb:

—  “shall” means that compliance with a requirement or a test is mandatory  for compliance with this  document;

—  “should” means that compliance with a requirement or a test is recommended but is not mandatory  for compliance with this document;

—  “may” is used to describe a permissible way to achieve compliance with a requirement or test.

An asterisk (*) as the first character of a title or at the beginning of a paragraph or table title indicates  that there is guidance or rationale related to that item in Annex AA.

The  attention  of  Member  Bodies  and  National  Committees  is  drawn  to  the  fact  that  equipment  manufacturers and testing organizations may need a transitional period following publication of a new,  amended  or  revised  ISO  or  IEC  publication  in  which  to  make  products  in  accordance  with  the  new  requirements and to equip themselves for conducting new or revised tests. It is the recommendation of  the committees that the content of this publication be adopted for implementation nationally not earlier  than 3  years  from  the date of publication  for equipment newly designed and not earlier  than 5  years  from the date of publication for equipment already in production.

© ISO 2017 – All rights reserved vii

BS EN ISO 80601-2-56:2017

BS EN ISO 80601-2-56:2017

INTERNATIONAL STANDARD ISO 80601-2-56:2017(E)

Medical electrical equipment —

Part 2‐56:

Particular requirements for basic safety and essential  performance of clinical thermometers for body temperature  measurement

201.1 * Scope, object and related standards

IEC 60601‐1:2005+A1:2012, Clause 1 applies, except as follows:

201.1.1 Scope

Replacement:

This  document  applies  to  the  BASIC  SAFETY  and  ESSENTIAL  PERFORMANCE  of  a  CLINICAL  THERMOMETER  in  combination  with its  ACCESSORIES,  hereafter  referred  to  as  ME EQUIPMENT.  This  document  specifies  the  general and  technical requirements  for electrical CLINICAL THERMOMETERS. This document applies  to all  electrical CLINICAL THERMOMETERS that are used for measuring the BODY TEMPERATURE of PATIENTS.

CLINICAL THERMOMETERS can be equipped with interfaces to accommodate secondary indicators, printing  equipment,  and  other  auxiliary  equipment  to  create  ME SYSTEMS.  This  document  does  not  apply  to  auxiliary equipment.

ME EQUIPMENT that measures a BODY TEMPERATURE is inside the scope of this document.

This document does not specify the requirements for screening thermographs intended to be used for  the individual non‐invasive human febrile temperature screening of groups of individual humans under  indoor environmental conditions, which are given in IEC 80601‐2‐59[4].

If a clause or subclause is specifically intended to be applicable to ME EQUIPMENT only, or to ME SYSTEMS only,  the  title and content  of  that clause  or  subclause will  say  so.  If  that is not  the case,  the clause  or  subclause applies both to ME EQUIPMENT and to ME SYSTEMS, as relevant.

HAZARDS inherent in the intended physiological function of ME EQUIPMENT or ME SYSTEMS within the scope  of  this  document  are  not  covered  by  specific  requirements  in  this  document  except  in  IEC 60601‐1:2005+A1:2012, 7.2.13 and 8.4.1.

NOTE  Additional information can be found in IEC 60601–1:2005+A1:2012, 4.2.

201.1.2 Object

Replacement:

The object of this particular document is to establish particular BASIC SAFETY and ESSENTIAL PERFORMANCE requirements for a CLINICAL THERMOMETER, as defined in 201.3.206, and its ACCESSORIES.

NOTE  ACCESSORIES  are  included  because  the  combination  of  the  CLINICAL  THERMOMETER  and  the  ACCESSORIES needs  to  be  safe  and  effective.  ACCESSORIES  can  have  a  significant  impact  on  the  BASIC  SAFETY  and  ESSENTIAL  PERFORMANCE of a CLINICAL THERMOMETER.

© ISO 2017 – All rights reserved 1

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.1.3 Collateral standards

Addition:

This  document  refers  to  those  applicable  collateral  standards  that  are  listed  in  IEC 60601‐1:2005+A1:2012, Clause 2, as well as 201.2 of this document.

IEC 60601‐1‐2,  IEC 60601‐1‐6,  IEC 60601‐1‐8,  IEC 60601‐1‐11  and  IEC 60601‐1‐12  apply  as modified  in Clauses 202, 206, 208, 211 and 212, respectively. IEC 60601‐1‐3[5] does not apply. All other published  collateral standards in the IEC 60601‐1 series apply as published.

201.1.4 Particular standards

Replacement:

In the IEC 60601 series, particular standards may modify, replace or delete requirements contained in  the general standard as appropriate for the particular ME EQUIPMENT under consideration, and may add  other BASIC SAFETY and ESSENTIAL PERFORMANCE requirements.

A requirement of a document takes priority over IEC 60601‐1 and its collateral standards.

For  brevity,  IEC 60601‐1:2005+A1:2012  is  referred  to  in  this  document  as  the  general  standard.  Collateral standards are referred to by their document number.

The numbering of sections, clauses and subclauses of this document corresponds to that of the general  standard  with  the  prefix  “201”  (e.g.  201.1  in  this  document  addresses  the  content  of  Clause 1  of  the  general standard) or applicable collateral standard with the prefix “20x” where x is the final digit(s) of  the collateral standard document number (e.g. 202.4 in this document addresses the content of Clause 4  of  the 60601‐1‐2 collateral  standard, 203.4 in  this document addresses  the content  of Clause 4  of  the  60601‐1‐3 collateral standard, etc.). The changes to the text of the general standard are specified by the  use of the following words:

“Replacement” means that the clause or subclause of the IEC 60601‐1 or applicable collateral standard  is replaced completely by the text of this particular document.

“Addition” means that the text of this document is additional to the requirements of the IEC 60601‐1 or  applicable collateral standard.

“Amendment” means that the clause or subclause of the IEC 60601‐1 or applicable collateral standard is  amended as indicated by the text of this document.

Subclauses or figures which are additional to those of the general standard are numbered starting from  201.101, Additional annexes are lettered AA, BB, etc., and additional items aa), bb), etc.

Subclauses or figures which are additional to those of a collateral standard are numbered starting from  20x,  where  “x”  is  the  number  of  the  collateral  standard,  e.g.  202  for  IEC 60601‐1‐2,  203  for  IEC 60601‐1‐3, etc.

The term “this document” is used to make reference to the IEC 60601‐1:2005+A1:2012, any applicable  collateral standards and this document taken together.

Where  there is no corresponding section, clause or subclause in  this document,  the section, clause or  subclause of  the  IEC 60601‐1 or applicable collateral standard, although possibly not relevant, applies  without  modification;  where  it  is  intended  that  any  part  of  the  IEC 60601‐1  or  applicable  collateral  standard,  although  possibly  relevant,  is  not  to  be  applied,  a  statement  to  that  effect  is  given  in  this  document.

2 © ISO 2017 – All rights reserved

201.2 Normative references

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

The  following  documents  are  referred  to  in  the  text  in  such  a  way  that  some  or  all  of  their  content  constitutes  requirements  of  this  document.  For  dated  references,  only  the  edition  cited  applies.  For  undated references, the latest edition of the referenced document (including any amendments) applies.

IEC 60601‐1:2005+A1:2012, Clause 2 applies, except as follows:

Replacement:

IEC 60601‐1‐2:2014, Medical electrical equipment — Part 1-2: General requirements for basic safety and  essential performance — Collateral Standard: Electromagnetic disturbances — Requirements and tests

IEC 60601‐1‐6:2010, Medical electrical equipment — Part 1-6: General requirements for basic safety and  essential performance — Collateral Standard: Usability +Amendment 1:2013

IEC 60601‐1‐8:2006, Medical electrical equipment — Part 1-8: General requirements for basic safety and  essential performance — Collateral Standard: General requirements, tests and guidance for alarm systems  in medical electrical equipment and medical electrical systems +Amendment 1:2012

Addition:

ISO 14155:2011, Clinical investigation of medical devices for human subjects — Good clinical practice

ISO 14937:2009, Sterilization of health care products — General requirements for characterization of a  sterilizing agent and the development, validation and routine control of a sterilization process for medical  devices

ISO 15223‐1:2016,  Medical devices — Symbols to be used with medical device labels, labelling and  information to be supplied — Part 1: General requirements

ISO 17664:2004, Sterilization of medical devices — Information to be provided by the manufacturer for  the processing of resterilizable medical devices

IEC 60601‐1:2005,  Medical electrical equipment — Part 1: General requirements for basic safety and  essential performance +Amendment 1:2012

IEC 60601‐1‐11:2015, Medical electrical equipment — Part 1-11: General requirements for basic safety  and essential performance — Collateral Standard: Requirements for medical electrical equipment and  medical electrical systems used in the home healthcare environment

IEC 60601‐1‐12:2014, Medical electrical equipment — Part 1-12: General requirements for basic safety  and essential performance — Collateral Standard: Requirements for medical electrical equipment and  medical electrical systems intended for use in the emergency medical services environment

IEC 62366‐1:2015, Medical devices — Part 1: Application of usability engineering to medical devices

© ISO 2017 – All rights reserved 3

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.3 Terms and definitions

For  the  purposes  of  this  document,  the  terms  and  definitions  given  in  IEC 60601‐1:2005+A1:2012,  IEC 60601‐1‐8:2006+A1:2012, IEC 62366‐1:2015 and the following apply.

ISO and IEC maintain terminological databases for use in standardization at the following addresses:  —  IEC Electropedia: available at http://www.electropedia.org/

—  ISO Online browsing platform: available at http://www.iso.org/obp

NOTE  An alphabetized index of defined terms is found beginning in Annex DD.

IEC 60601‐1:2005+A1:2012, Clause 3 applies, except as follows:

Additions:

201.3.201

* ADJUSTED MODE

OPERATING MODE where the OUTPUT TEMPERATURE is calculated by adjusting the signal from the input SENSOR

Note 1 to entry: For the purposes of this document, emissivity is considered a thermal or physiological property of  the MEASURING SITE, i.e. any CLINICAL THERMOMETER utilizing radiance that is dependent on emissivity is considered to  operate in an ADJUSTED MODE.

201.3.202

BLACKBODY

REFERENCE TEMPERATURE SOURCE of infrared radiation characterized by precisely known temperature and  having an effective emissivity close to one

201.3.203

BODY TEMPERATURE

all temperatures of the human body except SKIN TEMPERATURE

201.3.204

CLINICAL ACCURACY

closeness of agreement between the OUTPUT TEMPERATURE of a CLINICAL THERMOMETER and the true value  of the temperature of the REFERENCE BODY SITE that the CLINICAL THERMOMETER purports to represent

201.3.205

CLINICAL BIAS

Δcb

mean difference between OUTPUT TEMPERATURES of a CLINICAL THERMOMETER and a REFERENCE CLINICAL  THERMOMETER for the intended REFERENCE BODY SITE with specified LIMITS OF AGREEMENT when measured  from selected group of subjects

Note 1 to entry: LIMITS OF AGREEMENT can also be described as clinical uncertainty.

4 © ISO 2017 – All rights reserved

201.3.206

CLINICAL REPEATABILITY σr

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

pooled standard deviation (over a selected group of subjects) of changes in multiple OUTPUT  TEMPERATURES taken from the same subject at the same MEASURING SITE with the same CLINICAL  THERMOMETER by the same OPERATOR within a relatively short time

201.3.207

* CLINICAL THERMOMETER

ME EQUIPMENT used for measuring at the MEASURING SITE and indicating the temperature at the REFERENCE BODY SITE

Note 1 to entry: The MEASURING SITE can be the same as the REFERENCE BODY SITE.

201.3.208

* DIRECT MODE

OPERATING MODE of a CLINICAL THERMOMETER where the OUTPUT TEMPERATURE is an unadjusted  temperature that represents the temperature of the MEASURING SITE to which the PROBE is coupled

201.3.209

EXTENDED OUTPUT RANGE

OUTPUT TEMPERATURE range having one or both limits that are outside of the RATED OUTPUT RANGE 201.3.210

FLUID BATH

REFERENCE TEMPERATURE SOURCE containing fluid at a uniform temperature

EXAMPLE  Water, oil and air.

201.3.211

LABORATORY ACCURACY

closeness of agreement between the OUTPUT TEMPERATURE of a thermometer and the true value of the  measurand

Note 1 to entry: LABORATORY ACCURACY is sometimes referred to as maximum permissible error.  201.3.212

LIMITS OF AGREEMENT

LA

the magnitude of a potential disagreement between outputs of two CLINICAL THERMOMETERS equal to  double the standard deviation of OUTPUT TEMPERATURE differences when used on the same human  subject

Note 1 to entry: LIMITS OF AGREEMENT can also be described as clinical uncertainty.

201.3.213

MEASURING SITE

part of a PATIENT where the temperature is measured

EXAMPLE  Pulmonary artery, distal oesophagus, sublingual space in the mouth, rectum, ear canal, axilla  (armpit), forehead skin.

201.3.214

OPERATING MODE

state of a CLINICAL THERMOMETER that gives an OUTPUT TEMPERATURE of an intended REFERENCE BODY SITE

© ISO 2017 – All rights reserved 5

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.3.215

OUTPUT RANGE

span between the lowest and highest limits of OUTPUT TEMPERTURE where a CLINICAL THERMOMETER indicates OUTPUT TEMPERATURE within the specified characteristics of LABORATORY ACCURACY

201.3.216

OUTPUT TEMPERATURE

temperature indicated by a thermometer

Note 1 to entry: Methods of indication can include printed, spoken, displayed and remotely displayed.  201.3.217

PROBE

part of a CLINICAL THERMOMETER that provides a thermal coupling between the SENSOR and the PATIENT Note 1 to entry: Thermal coupling can be contact or non‐contact.

201.3.218

PROBE CABLE EXTENDER

cable that connects a CLINICAL THERMOMETER to a PROBE

Note 1 to entry: Not every CLINICAL THERMOMETER utilizes a PROBE CABLE EXTENDER.

Note 2 to entry: A PROBE CABLE EXTENDER can be an APPLIED PART.

201.3.219

PROBE COVER

disposable or reusable ACCESSORY of a CLINICAL THERMOMETER that provides a sanitary barrier between  the PROBE and the PATIENT

201.3.220

* REFERENCE BODY SITE

part of a PATIENT to which the OUTPUT TEMPERATURE refers

EXAMPLE  Pulmonary artery, distal oesophagus, sublingual space in the mouth, rectum, ear canal, axilla  (armpit), forehead skin.

201.3.221

REFERENCE CLINICAL THERMOMETER

RCT

CLINICAL THERMOMETER having established CLINICAL ACCURACY and LABORATORY ACCURACY, which is used for  CLINICAL ACCURACY VALIDATION of another CLINICAL THERMOMETER

201.3.222

REFERENCE TEMPERATURE SOURCE

source of a thermal energy whose temperature is measured by a REFERENCE THERMOMETER EXAMPLE  Blackbody, fluid bath.

201.3.223

REFERENCE THERMOMETER

equilibrium thermometer of a contact type for laboratory application whose calibration is traceable to a  national standard of temperature, with a specified accuracy and associated uncertainty

Note 1 to entry: An equilibrium thermometer can also be described as a zero‐heat‐flow thermometer.  EXAMPLE  Platinum resistance thermometer with calibration traceable to a national standard of temperature.

6 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.3.224

REPROCESSING

any activity, not specified in the ACCOMPANYING DOCUMENT, that renders a used product ready for re‐use  Note 1 to entry: The term “REPROCESSED” is used to designate the corresponding status.

Note 2 to entry: Such activities are often referred to as refinishing, restoring, recycling, refurbishing, repairing or  remanufacturing.

Note 3 to entry: Such activities can occur in healthcare facilities.

201.3.225

SENSOR

part of the CLINICAL THERMOMETER that converts thermal energy into an electrical signal  201.3.226

SKIN TEMPERATURE

temperature of the skin of the PATIENT at a point on which the sensing device intended to measure the temperature is placed

[SOURCE: IEC 60601‐2‐19:2009, 3.8.5, modified — replaced “infant” with “PATIENT” and “infant skin  temperature” with “the sensing device intended to measure the temperature is placed”]

201.3.227

TEST MODE

state of a CLINICAL THERMOMETER where the OUTPUT TEMPERATURE represents the temperature measured  by the SENSOR and is not adjusted for a REFERENCE BODY SITE or the rate of response of the SENSOR

Note 1 to entry: The TEST MODE can be used for the determination of the LABORATORY ACCURACY of the CLINICAL THERMOMETER.

Note 2 to entry: The TEST MODE can be the DIRECT MODE of the CLINICAL THERMOMETER.

201.3.228

VALIDATION

confirmation, through the provision of OBJECTIVE EVIDENCE, that the requirements for a specific INTENDED  USE or application have been fulfilled

Note 1 to entry: The term “VALIDATED” is used to designate the corresponding status.

Note 2 to entry: The use conditions for VALIDATION can be real or simulated.

[SOURCE: ISO 9000:2015, 3.8.13]

201.4 General requirements

IEC 60601‐1:2005+A1:2012, Clause 4 applies, except as follows:

201.4.2 RISK MANAGEMENT PROCESS for ME EQUIPMENT or ME SYSTEMS

Additional subclause:

201.4.2.101 Additional requirements for RISK MANAGEMENT PROCESS for ME EQUIPMENT or ME SYSTEMS

When  performing  the  RISK  MANAGEMENT  PROCESS  required  by  IEC 60601‐1:2005+A1:2012,  4.2,  the  analysis  shall  consider  the  RISKS  of  changing  environmental  conditions  for  the  CLINICAL  THERMOMETER and provide guidance regarding the RESIDUAL RISKS in the instruction for use.

NOTE  PORTABLE  CLINICAL  THERMOMETERS  can  undergo  changing  environmental  conditions  that  can  affect the  LABORATORY ACCURACY.

Compliance is checked by inspection of the instructions for use and the RISK MANAGEMENT FILE.

© ISO 2017 – All rights reserved 7

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.4.3 ESSENTIAL PERFORMANCE

Additional subclause:

201.4.3.101 * Additional requirements for ESSENTIAL PERFORMANCE

Additional ESSENTIAL PERFORMANCE requirements are found in the subclauses listed in Table 201.101.

Table 201.101 — Distributed ESSENTIAL PERFORMANCE requirements

| Requirement | Subclause |
| --- | --- |
| Accuracy of the CLINICAL THERMOMETER or at least one of the following: — generation of a TECHNICAL ALARM CONDITION; — not providing an OUTPUT TEMPERATURE. | 201.101.2 201.12.1.101 |

201.5 General requirements for testing of ME EQUIPMENT

IEC 60601‐1:2005+A1:2012, Clause 5 applies.

201.6 Classification of ME EQUIPMENT and ME SYSTEMS

IEC 60601‐1:2005, Clause 6 applies.

201.7 ME EQUIPMENT identification, marking and documents

IEC 60601‐1:2005+A1:2012, Clause 7 applies, except as follows:

201.7.2.1 Minimum requirements for marking on ME EQUIPMENT and interchangeable parts  Additional subclause:

201.7.2.1.101 Additional requirements for marking of the packaging

The packaging of a CLINICAL THERMOMETER and PROBE shall be marked with the following information:  a) MEASURING SITE and REFERENCE BODY site;

b) details  to enable  the OPERATOR  to identify  the mode of operation of  the CLINICAL  THERMOMETER and the contents of the packaging;

EXAMPLE  This package contains a predictive thermometer that estimates the PATIENT's temperature and 10 probe covers.

c) if sterile, the appropriate symbol from ISO 15223‐1:2016 (see Table 201.D.2.101, Symbols 3 to 8);

d) for  a  CLINICAL  THERMOMETER  or  PROBE  with  an  expiration  date  (use‐by  date),  ISO 15223‐1:2016, Symbol 5.1.4 (see Table 201.D.2.101, Symbol 2);

e) any special storage, handling or operating instructions.

For a specific MODEL OR TYPE REFERENCE, the indication of single use shall be consistent.  Compliance is checked by inspection.

201.7.2.2 Marking on the outside of ME EQUIPMENT or ME EQUIPMENT parts  Additional subclause:

8 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.7.2.101 Additional requirements for marking on the outside of ME EQUIPMENT or ME EQUIPMENT parts

A CLINICAL THERMOMETER shall have CLEARLY LEGIBLE markings with the following information:

a) the  symbol  “°C”  or  “°F”  adjacent  to  the  OUTPUT  TEMPERATURE,  if  not  indicated  at  the  display.  If switching  between  degrees  Fahrenheit  and  degrees  Celsius  is  possible,  the  respective  unit  of measure of the OUTPUT TEMPERATURE shall be indicated unambiguously;

b) intended MEASURING SITE;

c) that a new PROBE COVER shall be used prior to next measurement, if necessary, to maintain ESSENTIAL PERFORMANCE.

Compliance is checked by inspection.

201.7.4.3 Units of measurement

Additional subclause:

201.7.4.3.101 Additional requirements for unit of measure

A  CLINICAL  THERMOMETER  shall  express  the  temperature  in  either  degrees  Celsius,  °C,  or  degrees  Fahrenheit, °F, or both.

The CLINICAL THERMOMETER shall clearly indicate the unit of measure.

Compliance is checked by inspection and functional testing.

201.7.9 ACCOMPANYING DOCUMENT

201.7.9.1 Additional general requirements

Amendment (replace the first dash with):

—  name or trade name and address of

—  the MANUFACTURER, and

—  where  the  MANUFACTURER  does  not  have  an  address  within  the  locale,  an  authorized  representative within the locale

to which the RESPONSIBLE ORGANIZATION can refer;

201.7.9.2 Additional requirements for instructions for use

Additional subclause:

201.7.9.2.101 Instructions for use

The instructions for use shall include the following:

a) a summary of the USE SPECIFICATION as determined for IEC 62366‐1:2015;

b) the MEASURING SITE and REFERENCE BODY SITE of the CLINICAL THERMOMETER;

© ISO 2017 – All rights reserved 9

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

c) if  applicable,  the  recommended  minimum  measuring  time  and  minimum  time  between measurements for each intended MEASURING SITE;

d) the RATED OUTPUT RANGE for each intended REFERENCE BODY SITE;

e) the LABORATORY  ACCURACY in  the RATED OUTPUT  RANGE and, if equipped,  the LABORATORY  ACCURACY in the RATED EXTENDED OUTPUT RANGE;

f) for CLINICAL THERMOMETERS intended to be used with a PROBE COVER:

1) instructions for the application of a PROBE COVER;

2) information  about  the  behaviour  of  the  CLINICAL  THERMOMETER  when  used  without  the  PROBE COVER;

g) the information whether  the  CLINICAL  THERMOMETER is a DIRECT  MODE  or an ADJUSTED  MODE  CLINICAL  THERMOMETER;

h) if applicable, instructions for selection and replacement of the INTERNAL ELECTRICAL POWER SOURCE; EXAMPLE 1  Battery replacement.

i) details of the nature and frequency of any maintenance and/or calibration needed to ensure that the CLINICAL THERMOMETER operates properly and safely;

j) information concerning the disposal of the CLINICAL THERMOMETER and its components; EXAMPLE 2  Battery or PROBE COVER disposal.

k) if the CLINICAL THERMOMETER or its parts are intended  for single use, information on characteristics and technical factors known to the MANUFACTURER that could pose a RISK if the CLINICAL THERMOMETER or its parts would be re‐used.

Compliance is checked by inspection.

Additional subclause:

201.7.9.101 Additional requirements for ACCOMPANYING DOCUMENT

Unless  the  CLINICAL  THERMOMETER  is  equipped  with  a  TEST  MODE  or  DIRECT  MODE,  the  ACCOMPANYING  DOCUMENT  shall  include  the  correction  method  to  derive  unadjusted  temperatures  from  OUTPUT  TEMPERATURE measured in the ADJUSTED MODE.

Compliance is checked by inspection.

201.8 Protection against electrical HAZARDS from ME EQUIPMENT

IEC 60601‐1:2005+A1:2012, Clause 8 applies.

201.9 Protection against mechanical HAZARDS of ME EQUIPMENT and ME SYSTEMS IEC 60601‐1:2005+A1:2012, Clause 9 applies.

201.10 Protection against unwanted and excessive radiation HAZARDS

IEC 60601‐1:2005+A1:2012, Clause 10 applies.

10 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.11 Protection against excessive temperatures and other HAZARDS

IEC 60601‐1:2005+A1:2012, Clause 11 applies, except as follows:

Additional subclause:

201.11.6.6 * Cleaning and disinfection of ME EQUIPMENT and ME SYSTEM

Amendment (add additional requirement as new first paragraph):

The  surfaces  of  the CLINICAL  THERMOMETER,  PROBE and ACCESSORIES  that can become contaminated with  body  fluids during NORMAL CONDITION or SINGLE FAULT CONDITION shall be designed  to allow  for cleaning  and  disinfection  or  cleaning  and  sterilization  (additional  requirements  are  found  in  IEC 60601‐1:2005+A1:2012,  11.6.7  and  IEC 60601‐1‐11:2015,  Clause 8).  The  CLINICAL  THERMOMETER may be dismantled during the decontamination PROCESS. ACCESSORIES not intended for re‐use are exempt  from this requirement.

Amendment (add additional requirement and replace the compliance test):

CLINICAL  THERMOMETER  and  PROBE  ENCLOSURES  shall  be  designed  to  allow  for  surface  cleaning  and  disinfection to reduce to acceptable levels the RISK of infection of OPERATORS, bystanders, or the PATIENT.

Instructions  for  processing  and  REPROCESSING  the  CLINICAL  THERMOMETER,  PROBE  and  ACCESSORIES  shall  comply with ISO 17664:2004 and ISO 14937:2009 and shall be disclosed in the instructions for use.

NOTE 1  ISO 14159[8] provides guidance for the design of ENCLOSURES.

Check compliance by inspection of the RISK MANAGEMENT FILE. When compliance with this document could  be affected by the cleaning or the disinfecting of the CLINICAL THERMOMETER, PROBE or its parts or  ACCESSORIES, clean and disinfect them 100 times in accordance with the methods indicated in the  instruction for use, including any cooling or drying period. After these PROCEDURES, confirm that BASIC  SAFETY and ESSENTIAL PERFORMANCE are maintained. Confirm that the MANUFACTURER has evaluated the  effects of multiple PROCESS cycles and the effectiveness of those cycles.

NOTE 2  Additional information regarding the order of test is found in 211.10.1.1 and 212.10.1.1.  201.11.7 Biocompatibility of ME EQUIPMENT and ME SYSTEMS

Amendment (add after existing text):

The ACCESSIBLE PARTS of a CLINICAL THERMOMETER, PROBE, its parts or ACCESSORIES that contain phthalates  which are classified as carcinogenic, mutagenic or toxic to reproduction shall be marked as containing  phthalates  on  the  device  itself  or  on  the  packaging.  The  symbol  of  EN 15986:2011  (see  Table 201.D.2.101, Symbol 10) may  be used.  If  the INTENDED  USE  of a CLINICAL  THERMOMETER,  PROBE, its  parts  or  ACCESSORIES  includes  treatment  of  children  or  treatment  of  pregnant  or  nursing  women,  a  specific justification  for the use of  these phthalates shall be included in  the RISK MANAGEMENT FILE. The  instructions  for  use  of  a  CLINICAL  THERMOMETER,  PROBE,  its  parts  or  ACCESSORIES  that  contain  such  phthalates  shall  contain information  on  RESIDUAL  RISKS  for  these  PATIENT  groups  and, if  applicable,  on  appropriate precautionary measures.

Amendment (replace compliance check):

Check compliance by inspection of the instructions for use and inspection of the information provided by  the MANUFACTURER for identification of the presence of substances which are carcinogenic, mutagenic or  toxic to reproduction and justification for their use.

© ISO 2017 – All rights reserved 11

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.12 Accuracy of controls and instruments and protection against hazardous outputs  IEC 60601‐1:2005+A1:2012, Clause 12 applies, except as follows:

201.12.1 Accuracy of controls and instruments

Additional subclause:

201.12.1.101 Additional requirements for accuracy of controls and instruments

When  the  CLINICAL  THERMOMETER  is  not  capable  of  indicating  a  temperature  within  the  LABORATORY  ACCURACY, it shall provide a TECHNICAL ALARM CONDITION or it shall not provide an OUTPUT TEMPERATURE.

NOTE 1  Possible causes include the following:

a) low voltage of the internal electrical power source;

b) output temperature outside the rated output range or rated extended output range.

The  OUTPUT  TEMPERATURE  of  CLINICAL  THERMOMETERS  should  cover  the  minimum  RATED  OUTPUT  RANGE from 34,0 °C to 43,0 °C.

NOTE 2  In some applications, a wider RATED OUTPUT RANGE can be needed.

NOTE 3  In  some  applications,  a  narrower  RATED  OUTPUT  RANGE  can  be  needed  (e.g.  ovulation  CLINICAL THERMOMETER).

Compliance is checked by inspection and functional testing.

201.12.2 Usability

Additional subclause:

201.12.2.101 * Additional requirements for USABILITY of ME EQUIPMENT

In  addition  to  the  requirements  of  IEC 60601‐1:2005+A1:2012,  12.2  for  CLINICAL  THERMOMETERS intended for use in HOME HEALTHCARE ENVIRONMENT, the display of OUTPUT TEMPERATURE shall be at least  4 mm high or optically magnified so as to appear that height.

For CLINICAL THERMOMETERS with a segmented indication display, a  functional test of all segments shall  be performed after activation.

Compliance is checked by inspection and inspection of the USABILITY ENGINEERING FILE. 201.13 HAZARDOUS SITUATIONS and fault conditions

IEC 60601‐1:2005+A1:2012, Clause 13 applies.

201.14 PROGRAMMABLE ELECTRICAL MEDICAL SYSTEMS (PEMS)

IEC 60601‐1:2005+A1:2012, Clause 14 applies.

201.15 Construction of ME EQUIPMENT

IEC 60601‐1:2005+A1:2012, Clause 15 applies.

12 © ISO 2017 – All rights reserved

201.16 ME SYSTEMS

IEC 60601‐1:2005+A1:2012, Clause 16 applies.

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

201.17 Electromagnetic compatibility of ME EQUIPMENT and ME SYSTEMS

IEC 60601‐1:2005, Clause 17 applies.

Additional subclauses:

201.101 Laboratory performance requirements

201.101.1 * General test requirements

Laboratory performance shall be assessed in DIRECT MODE or in TEST MODE under the same conditions. If  a  TEST  MODE  or  DIRECT  MODE  is  not  available,  use  the  correction  method  to  derive  unadjusted  temperatures from the OUTPUT TEMPERATURE in accordance with the ACCOMPANYING DOCUMENT.

A  CLINICAL  THERMOMETER intended  for  use with  a  PROBE  COVER  shall  be  tested  together with  the  PROBE  COVER  as  indicated  in  the  instructions  for  use.  A  new  PROBE  COVER  shall  be  used  for  each  OUTPUT  TEMPERATURE measurement.

NOTE  A CLINICAL THERMOMETER OPERATING MODE whose accuracy cannot be VERIFIED by utilizing a FLUID BATH is  also required to be evaluated according to 201.102.

Compliance is checked by inspection and inspection of the ACCOMPANYING DOCUMENT. 201.101.2 * LABORATORY ACCURACY

The LABORATORY ACCURACY within the RATED OUTPUT RANGE in NORMAL USE shall be within ±0,3 °C.

The  LABORATORY  ACCURACY,  within  the  RATED  EXTENDED  OUTPUT  RANGE  in  NORMAL  USE,  shall  be  within  ±0,4 °C unless CLINICAL  THERMOMETER indicates  that  the measured  temperature is outside of  the RATED  OUTPUT RANGE.

If the PROBE is separable from the CLINICAL THERMOMETER, they may be tested separately.

NOTE  For  some  applications,  improved  LABORATORY  ACCURACY  can  be  needed  (e.g.  ovulation  CLINICAL THERMOMETER).

Compliance is checked with the following test.

a) * Utilize a REFERENCE TEMPERATURE SOURCE, which is either a FLUID BATH or a BLACKBODY radiator as described in Annex BB. Confirm that the temperature supplied by the REFERENCE TEMPERATURE SOURCE has an expanded measurement uncertainty (covering factor k = 2) of not greater than 0,07 °C.

b) Place the CLINICAL THERMOMETER into a climatic chamber and set the ambient temperature and humidity to approximately the mid-range of the RATED temperature and humidity as indicated in the ACCOMPANYING DOCUMENT.

c) Set the temperature of the REFERENCE TEMPERATURE SOURCE to approximately the midpoint of the RATED  OUTPUT RANGE. Stabilize the CLINICAL THERMOMETER at given conditions of ambient temperature and humidity for a minimum of 30 min, or longer if indicated in the ACCOMPANYING DOCUMENT.

d) Measure the temperature of the REFERENCE TEMPERATURE SOURCE with both the CLINICAL THERMOMETER and a REFERENCE THERMOMETER. Document the results.

© ISO 2017 – All rights reserved 13

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

e) Repeat c) and d), once setting the temperature of the REFERENCE TEMPERATURE SOURCE to within 1 °C of the upper limit and once within 1 °C of the lower limit of the RATED OUTPUT RANGE.

f) If the CLINICAL THERMOMETER is provided with an EXTENDED OUTPUT RANGE, repeat e),  once setting the temperature of the REFERENCE TEMPERATURE SOURCE to within 0,5 °C of the upper limit and once within 0,5 °C of the lower limit of the RATED EXTENDED OUTPUT RANGE.

g) Repeat c) through f) at four points, the combinations of upper and lower limits of the RATED environmental temperature and humidity ranges as indicated in the ACCOMPANYING DOCUMENT.

h) Calculate the measurement error, e, for each individual OUTPUT TEMPERATURE measurement using Formula (1).

i) Confirm that the measurement error meets the requirement.

The error of an individual OUTPUT TEMPERATURE measurement is indicated by Formula (1). e = TDUT − TREF (1)  where

TDUT is the individual OUTPUT TEMPERATURE of the CLINICAL THERMOMETER under test (DUT) when  measuring the REFERENCE TEMPERATURE SOURCE;

TREF is the corresponding OUTPUT TEMPERATURE of the REFERENCE TEMPERATURE SOURCE measured by  the REFERENCE THERMOMETER.

201.101.3 * Time response for a DIRECT MODE CLINICAL THERMOMETER that is not an ADJUSTED MODE  CLINICAL THERMOMETER

The  transient  response  for a DIRECT  MODE  CLINICAL  THERMOMETER  that is  not an ADJUSTED MODE  CLINICAL  THERMOMETER shall be characterized and disclosed in the instructions for use.

Compliance is checked by inspection of the instructions for use and with the following test.

a) Utilize two REFERENCE TEMPERATURE SOURCES, which are FLUID BATHS, BLACKBODIES or specially designed thermal sources as described in Annex BB.

b) Set the temperature of the first REFERENCE TEMPERATURE SOURCE to approximately the middle of the RATED OUTPUT RANGE. Set the temperature of the second REFERENCE TEMPERATURE SOURCE to approximately 2 °C higher than the first REFERENCE TEMPERATURE SOURCE. Confirm that the temperature of the second REFERENCE TEMPERATURE SOURCE is within the RATED OUTPUT RANGE of the CLINICAL  THERMOMETER.

c) For 5 min, or longer if necessary to achieve thermal equilibrium, thermally couple the PROBE of the CLINICAL THERMOMETER to the first REFERENCE TEMPERATURE SOURCE. Note the time and immediately move the PROBE to the second REFERENCE TEMPERATURE SOURCE.

d) Continuously monitor the OUTPUT TEMPERATURE until the new temperature reading remains within the LABORATORY ACCURACY limits of the second REFERENCE TEMPERATURE SOURCE and is within the LABORATORY ACCURACY of the CLINICAL THERMOMETER. Note the time.

e) Note the difference between the two points in time and document the heating transient time.

f) Set the temperature of the second REFERENCE TEMPERATURE SOURCE to approximately 2 °C lower than the first REFERENCE TEMPERATURE SOURCE. Confirm that the temperature of the second REFERENCE  TEMPERATURE SOURCE is within the RATED OUTPUT RANGE of the CLINICAL THERMOMETER.

14 © ISO 2017 – All rights reserved

g) Repeat c) and d).

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

h) Note the difference between the two times and document the cooling transient time.

i) Confirm that the transient time for both heating and cooling is less than the recommended minimum measuring time indicated in the ACCOMPANYING DOCUMENT.

201.102 * CLINICAL ACCURACY VALIDATION

201.102.1 Method

ADJUSTED MODE CLINICAL THERMOMETERS shall be VALIDATED  for CLINICAL ACCURACY in each ADJUSTED MODE.  The  results  of  CLINICAL  ACCURACY  VALIDATION  shall  be  disclosed  in  the  ACCOMPANYING  DOCUMENT.  This  disclosure  shall  include:  the  CLINICAL  BIAS,  Δcb,  with  its  LIMITS  OF  AGREEMENT,  LA,  the  CLINICAL REPEATABILITY, σr, the REFERENCE BODY SITE and MEASURING SITE for each OPERATING MODE.

NOTE 1  CLINICAL BIAS is  the mean difference between OUTPUT TEMPERATURES of  the CLINICAL THERMOMETER under  test  (DUT) and  the RCT  for a specific REFERENCE BODY SITE when measured  from a selected group of subjects. The  CLINICAL BIAS for each OPERATING MODE of the DUT defines closeness between the DUT OUTPUT TEMPERATURE and that  of the RCT.

NOTE 2  Any CLINICAL THERMOMETER OPERATING MODE whose accuracy cannot be VERIFIED by utilizing a FLUID BATH is considered an ADJUSTED MODE CLINICAL THERMOMETER.

NOTE 3  The CLINICAL THERMOMETER can have more than one OPERATING MODE with different CLINICAL ACCURACIES.  The CLINICAL ACCURACY VALIDATION shall be conducted in accordance with ISO 14155:2011.

NOTE 4  This  document  specifies  that  during  a  CLINICAL  ACCURACY  VALIDATION  of  a  CLINICAL  THERMOMETER,  characteristics  that  represent  CLINICAL  ACCURACY  (i.e.  the  CLINICAL  BIAS  with  its  LIMITS  OF  AGREEMENT  and  CLINICAL  REPEATABILITY) are evaluated. These characteristics are evaluated from the same data set obtained with a CLINICAL THERMOMETER  under  test  (DUT)  and  a  REFERENCE  CLINICAL  THERMOMETER  (RCT),  both  indicating  the  OUTPUT TEMPERATURE of the same REFERENCE BODY SITE.

A  CLINICAL  THERMOMETER intended  for  use with  a  PROBE  COVER  shall  be  tested  together with  the  PROBE  COVER  as  indicated  in  the  instructions  for  use.  A  new  PROBE  COVER  shall  be  used  for  each  OUTPUT  TEMPERATURE measurement.

Consider compliance with the requirements of this subclause to exist when the requirements of 201.101.2,  201.102.3, 201.102.4 and 201.102.5 have been fulfilled, by inspection of the instructions for use and with  the following test.

a) Obtain temperatures with the DUT in the OPERATING MODE being evaluated and with the PATIENT population indicated in the instructions for use according to 201.101.2. Before and after the tests, confirm the LABORATORY ACCURACY of the RCT with a REFERENCE TEMPERATURE SOURCE.

b) Take OUTPUT TEMPERATURES from the DUT and the RCT concurrently, or sequentially, when using the same MEASURING SITE, using the measuring time as indicated in the instructions for use. Use the RCT according to its instructions for use.

c) Perform at least three consecutive measurement PROCEDURES, as indicated in the instructions for use, with the DUT and at least one temperature with an RCT for each subject. Wait between individual measurements as indicated in the instructions for use. See also 201.101.1. For febrile subjects less than 5 years of age, only one measurement need be taken.

NOTE 5  Care should be taken to ensure that the actual temperatures of the MEASURING SITES of the DUT and RCT remain stable during a series of measurements.

© ISO 2017 – All rights reserved 15

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

d) Repeat b) and c) for each subject in the study. Take care to exclude subjects with medical conditions which might skew CLINICAL ACCURACY VALIDATION results such as inflammation at the MEASURING SITE.

EXAMPLE 1  Barbiturates, thyroid preparations, antipsychotics, recent immunizations. Take care to exclude subjects who have taken antipyretics in the preceding 120 min.

NOTE 6  Studies  show  that  PATIENT  temperature  is  reasonably  stable  120 min  after  taking  common antipyretics[9][10].

EXAMPLE 2  Aspirin, acetaminophen, ibuprofen.

201.102.2 * Human subject population requirements

The total number of febrile subjects shall be not less than 30 % and not greater than 50 % of all subjects  in the selected age group.

NOTE 1  The  first  temperature  out  of  three measurements is used in calculation  of CLINICAL  BIAS and  LIMITS  OF AGREEMENT. All three OUTPUT TEMPERATURE measurements are used in calculation of the CLINICAL REPEATABILITY.

NOTE 2  For the purpose of CLINICAL ACCURACY VALIDATION, a febrile subject is defined as a subject having  —  an elevated core or rectal temperature of 38,0 °C (100,4 °F) or higher as measured by an RCT,  —  an elevated sublingual temperature of 37,5 °C (99,5 °F) or higher as measured by an RCT, or  —  an elevated axillary temperature of 37,2 °C (99,0 °F) or higher as measured by an RCT.

CLINICAL ACCURACY VALIDATION shall be carried out on all age groups indicated in the instructions for use  as  indicated  in  the  USE  SPECIFICATION.  The  minimum  number  of  subjects  in  the  CLINICAL  ACCURACY  VALIDATION  shall  be  not less  than  105.  The  number  of  subjects in  each  age  group  shall  be  sufficiently  large  to  minimize  the  effect  of  random  components  of  measurement  error.  The  age  groups  shall  be  developed  with  the  guidelines  of  Table 201.102.  Subjects  may  be PATIENTS.  The  minimum  number  of  subjects in an age group shall be at least 35, and at least 15 in subgroup A1 or A2 if subgroup A is not  explicitly excluded in the instructions for use.

Table 201.102 — Subject age groups

| Age group | Age[11] |
| --- | --- |
| A1 | 0 up to 3 months |
| A2 | 3 months up to 1 year |
| B | older than 1 year and younger than 5 years |
| C | older than 5 years |

16 © ISO 2017 – All rights reserved

201.102.3 * CLINICAL BIAS calculation

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

To evaluate the CLINICAL BIAS for the OPERATING MODE being evaluated, use the first CLINICAL THERMOMETER under test (DUT) OUTPUT TEMPERATURE out of three and the corresponding RCT OUTPUT TEMPERATURE from  each  subject in  a  test  group.  The  CLINICAL  BIAS  for  the  specified  REFERENCE  BODY  SITE  and  age  group is  calculated from Formula (2).

n

∑ −

( ) DUT, , RCT

T T

i i

i

Δcb =

where

=

1

n

(2)

i is the index number for an individual subject;

n is the total number of subjects per MEASURING SITE and age group;

TDUT,i  is the ith observed OUTPUT TEMPERATURE from the DUT;

TRCT,i  is the ith observed OUTPUT TEMPERATURE from the RCT.

201.102.4 * LIMITS OF AGREEMENT calculation

To calculate the LIMITS OF AGREEMENT, LA, use Formula (3).

2 Δ ×σ (3) LA = cb

where

cb Δ σ is calculated using Formula (4).

To  calculate  the  deviation, cb Δ σ ,  use  the  first  measurement  of  the  CLINICAL  THERMOMETER  under  test (DUT) OUTPUT TEMPERATURE out of three measurements and the corresponding RCT OUTPUT TEMPERATURE of each subject for the OPERATING MODE being evaluated using Formula (4).

n

∑

( ) RCT2

Δ

⎡ ⎤ − − ⎢ ⎥ ⎣ ⎦

T T

DUT cb , ,

i i

i

cb Δ σ  =

where

=

1

1

n

−

(4)

i is the index number for an individual subject;

n is the total number of subjects per MEASURING SITE and age group;  TDUT,i  is the ith OUTPUT TEMPERATURE indicated by the DUT;  TRCT,i  is the ith OUTPUT TEMPERATURE indicated by the RCT;  Δcb is the CLINICAL BIAS as calculated in Formula (2).

© ISO 2017 – All rights reserved 17

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

201.102.5 * CLINICAL REPEATABILITY calculation

An  ADJUSTED  MODE  CLINICAL  THERMOMETER  that makes  continuous  estimates  of  the  REFERENCE  BODY  SITE temperature shall be exempt from the requirements of this subclause.

CLINICAL REPEATABILITY,  for a particular OPERATING MODE, is determined  for  the subject population of all  age groups given in Table 201.102 combined. Febrile subjects less than 5 years of age may be excluded.

CLINICAL REPEATABILITY is calculated by a pooled standard deviation of triplicate measurements over the  entire population of subjects. First, calculate the standard deviation, σj, of the three OUTPUT TEMPERATURE measurements (T1,j, T2,j and T3,j) for each subject j using Formula (5).

m

∑2

( )

T T

−

, ,

i j

DUT DUT

i

σ =

1

m(5)

j

where

=

−

1

DUT,i T is the ith OUTPUT TEMPERATURE (e.g. 1, 2 or 3) indicated by the DUT;

DUT, j T is the average of the OUTPUT TEMPERATURES on subject j;

m is the number of OUTPUT TEMPERATURE measurements on the subject.

NOTE  m is typically equal to 3.

Then  calculate  a  pooled  standard  deviation  (the  CLINICAL  REPEATABILITY),  σr,  for  all  subjects  using  Formula (6).

σ + ++ + + σσ σ 22 2 2

σr =

where

1 2 ... .... j N

N (6)

N is the total number of subjects of all age groups in a study.

201.103 * PROBES, PROBE CABLE EXTENDERS and PROBE COVERS

201.103.1 General

All  PROBES,  PROBE  CABLE  EXTENDERS  and  PROBE  COVERS  shall  comply  with  the  requirements  of  this  document, whether they are produced by the MANUFACTURER of the CLINICAL THERMOMETER or by another  entity (“third party manufacturer” or healthcare provider) or are REPROCESSED.

MANUFACTURERS of a PROBE, PROBE CABLE EXTENDER and PROBE COVER, including a REPROCESSED PROBE, PROBE  CABLE  EXTENDER  and  PROBE  COVER,  shall  conduct  tests  to  VERIFY  that  all  CLINICAL  THERMOMETER specifications  are  met  with  each  MODEL  OR  TYPE  REFERENCE  of  CLINICAL  THERMOMETER  with  which  the  PROBE,  PROBE  CABLE  EXTENDER  or  PROBE  COVER  is  intended  to  be  used.  The  ACCOMPANYING  DOCUMENT  of  PROBES,  PROBE  CABLE  EXTENDERS  and  PROBE  COVERS,  including  those  that  are  REPROCESSED,  shall  list  all  CLINICAL THERMOMETERS with which compatibility is claimed.

It is the responsibility of the MANUFACTURER of a PROBE, PROBE CABLE EXTENDER and PROBE COVER, including  a REPROCESSED PROBE, PROBE CABLE EXTENDER and PROBE COVER, to have their PROCESSES VALIDATED to VERIFY that any new or REPROCESSED product complies with the requirements of this document.

18 © ISO 2017 – All rights reserved

Compliance is checked by the tests of this document. 201.103.2 Labelling

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

The MODEL OR TYPE REFERENCE of at least one CLINICAL THERMOMETER shall be included in the ACCOMPANYING  DOCUMENT provided with each PROBE, PROBE CABLE EXTENDER and PROBE COVER, compliant with 201.103.1.

Statements shall be included in  the ACCOMPANYING  DOCUMENT  of each PROBE, PROBE  CABLE  EXTENDER and  PROBE COVER to the effect that

a) they are designed for use with specific thermometer or monitoring equipment,

b) the  operator  is  responsible  for  checking  the  compatibility  of  the  thermometer  or  monitoring equipment, probe, probe cable extender and probe cover before use, and

c) incompatible components can result in degraded performance.

Additional information is found in 201.103.1.

Compliance is checked by inspection of the ACCOMPANYING DOCUMENT.

202 Electromagnetic disturbances — Requirements and tests

IEC 60601‐1‐2:2014 applies except as follows:

202.4.3.1 Configurations

Amendment (add after the last dash of 4.3.1):

—  if  applicable,  attachment  of  ACCESSORIES  as  necessary  to  achieve  the  BASIC  SAFETY  and  ESSENTIAL  PERFORMANCE of the CLINICAL THERMOMETER;

202.5.2.2.1 Requirements applicable to all ME EQUIPMENT and ME SYSTEMS

Amendment [add note to list element b)]:

NOTE  The requirements of this document are not considered deviations or allowances.  Addition:

202.8.1.101 Additional general requirements

The  following  degradations,  if  associated  with  BASIC  SAFETY  or  ESSENTIAL  PERFORMANCE,  shall  not  be  allowed:

—  component failures;

—  changes in programmable parameters or settings;

—  reset to default settings;

—  change of OPERATING MODE;

—  change in LABORATORY ACCURACY at any point in the RATED OUTPUT RANGE and in the RATED EXTENDED  OUTPUT  RANGE  greater  than  permitted  in  201.101.2  without  the  generation  of  either  a  TECHNICAL  ALARM CONDITION or an indication of abnormal operation.

© ISO 2017 – All rights reserved 19

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

The CLINICAL THERMOMETER may exhibit temporary degradation of performance (e.g. deviation from the  performance  indicated  in  the  instructions  for  use  during  IMMUNITY  testing),  which  does  not  require  OPERATOR intervention and that does not affect BASIC SAFETY or ESSENTIAL PERFORMANCE.

In  the  event  of  disruption  during  IMMUNITY  tests,  the  CLINICAL  THERMOMETER  shall  recover  from  any  disruption within 30 s.

206 Usability

IEC 60601‐1‐6:2010+A1:2013 applies except as follows:

For a CLINICAL THERMOMETER, the following shall be considered PRIMARY OPERATING FUNCTIONS:  a) observing the OUTPUT TEMPERATURE;

b) properly positioning CLINICAL THERMOMETER or PROBE at the MEASURING SITE;

c) starting the CLINICAL THERMOMETER from power off;

d) turning off the CLINICAL THERMOMETER;

e) performing a basic pre‐use functional check of the CLINICAL THERMOMETER.

The following functions, if available, also shall be considered PRIMARY OPERATING FUNCTIONS:  f) performing a basic pre‐use functional check of the ALARM SIGNALS;

g) configuring  the  ACCESSORIES  including  connection  of  the  detachable  parts  to  the  CLINICAL THERMOMETER;

h) REPROCESSING the ACCESSORIES;

i) setting the OPERATOR‐adjustable controls:

—  switching between MEASURING SITES;

—  setting CLINICAL THERMOMETER control parameters;

—  setting ALARM LIMITS;

—  inactivating ALARM SIGNALS.

208 General requirements, tests and guidance for alarm systems in medical electrical equipment  and medical electrical systems

IEC 60601‐1‐8:2006+A1:2012 applies except as follows:

208.6.1.2 Determination of ALARM CONDITIONS and assignment of priority  Amendment (add prior to the compliance check):

A CLINICAL THERMOMETER need not have a PHYSIOLOGICAL ALARM CONDITION.

20 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

211 Requirements for medical electrical equipment and medical electrical systems used in the  home healthcare environment

IEC 60601‐1‐11:2015 applies except as follows:

211.4.2.3.1 Continuous operating conditions

Amendment (replace the first dash with):

—  a temperature range of at least +15 °C to +40 °C;

211.10.1.1 General requirements for mechanical strength

Amendment (add as the first paragraph):

The tests of IEC 60601‐1‐11:2015, Clause 10 and IEC 60601‐1:2005+A1:2012, 15.3 shall be performed  on  the  same  sample  of  the  CLINICAL  THERMOMETER  after  the  cleaning  and  disinfection  PROCEDURES  of  201.11.6.6  of  this  document  have  been  performed  unless  there  are  no  cleaning  and  disinfection  PROCEDURES  specified  in  the  instructions  for  use.  If  more  than  one  PROCEDURE  is  specified  in  the  instructions for use, each PROCEDURE shall be so tested. A separate sample of the CLINICAL THERMOMETER may be used for each specified PROCEDURE.

212 Requirements for medical electrical equipment and medical electrical systems intended for  use in the emergency medical services environment

IEC 60601‐1‐12:2014 applies except as follows:

212.10.1.1 General requirements for mechanical strength

Amendment (add as the first paragraph):

The tests of IEC 60601‐1‐12:2014, Clause 10 and IEC 60601‐1:2005+A1:2012, 15.3 shall be performed  on  the  same  sample  of  the  CLINICAL  THERMOMETER  after  the  cleaning  and  disinfection  PROCEDURES  of  201.11.6.6  of  this  document  have  been  performed  unless  there  are  no  cleaning  and  disinfection  PROCEDURES  specified  in  the  instructions  for  use.  If  more  than  one  PROCEDURE  is  specified  in  the  instructions for use, each PROCEDURE shall be so tested. A separate sample of the CLINICAL THERMOMETER may be used for each specified PROCEDURE.

Annexes

IEC 60601‐1:2005+A1:2012, annexes apply, except as follows:

© ISO 2017 – All rights reserved 21

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Annex C

(informative)

Guide to marking and labelling requirements for ME EQUIPMENT and ME SYSTEMS

Annex C of the general standard applies, except as follows:

201.C.1 Marking on the outside of ME EQUIPMENT, ME SYSTEMS, or their parts

Addition:

201.C.1.101Marking on the outside of a CLINICAL THERMOMETER or its parts

Additional requirements for marking on the outside of a CLINICAL THERMOMETER or its parts are found in  Table 201.C.101.

Table 201.C.101 — Marking on the outside of a CLINICAL THERMOMETER or its parts

| Description of marking | Subclause |
| --- | --- |
| Containing phthalates, if applicable | 201.11.7 |
| Intended MEASURING SITE | 201.7.2.101 b) |
| On the packaging, any special storage, handing or operating instructions | 201.7.2.1.101 e) |
| On the packaging, appropriate sterility symbol, if applicable | 201.7.2.1.101 c) |
| On the packaging, containing phthalates, if applicable | 201.11.7 |
| On the packaging, expiration date, if applicable | 201.7.2.1.101 d) |
| On the packaging, MEASURING SITE | 201.7.2.1.101 a) |
| On the packaging, REFERENCE BODY SITE | 201.7.2.1.101 a) |
| On the packaging, mode of operation of CLINICAL THERMOMETER | 201.7.2.1.101 b) |
| On the packaging, contents of package | 201.7.2.1.101 b) |
| Symbol “°C” or “°F”, as appropriate | 201.7.2.101 a) |
| That a new PROBE COVER shall be used prior to the next measurement, if applicable | 201.7.2.101 c) |
| Clearly indicate the units of measurement | 201.7.4.3.101 |

201.C.4 ACCOMPANYING DOCUMENTS, general

Addition:

201.C.4.101 ACCOMPANYING DOCUMENTS, general, of a CLINICAL THERMOMETER

Additional requirements for ACCOMPANYING DOCUMENTS, general, of a CLINICAL THERMOMETER are found in  Table 201.C.102.

22 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Table 201.C.102 — Accompanying documents, general, of a clinical thermometer

| Description of marking | Subclause |
| --- | --- |
| Correction method to derive the unadjusted temperature from the OUTPUT TEMPERATURE measured in the ADJUSTED MODE, if applicable | 201.7.9.101 |
| Correction method to derive the unadjusted temperature from the OUTPUT TEMPERATURE, if applicable | 201.101.1 |
| For a PROBE, PROBE CABLE EXTENDER and PROBE COVER, all CLINICAL THERMOMETERS with which compatibility is claimed | 201.103.1 |
| For a PROBE, PROBE CABLE EXTENDER and PROBE COVER, a statement to the effect that it is designed for use with specific thermometer or monitoring equipment | 201.103.2 a) |
| For a PROBE, PROBE CABLE EXTENDER and PROBE COVER, a statement to the effect that the OPERATOR is responsible for checking the compatibility of the thermometer or monitoring equipment, probe, probe cable extender and probe cover before use | 201.103.2 b) |
| For a PROBE, PROBE CABLE EXTENDER and PROBE COVER, a statement to the effect that incompatible components can result in degraded performance | 201.103.2 c) |
| For a PROBE, PROBE CABLE EXTENDER and PROBE COVER, the MODEL OR TYPE REFERENCE of at least one CLINICAL THERMOMETER | 201.103.2 |
| For an ADJUSTED MODE CLINICAL THERMOMETER, the results of the CLINICAL ACCURACY VALIDATION | 201.102.1 |
| Name and address of the MANUFACTURER and where the MANUFACTURER does not have an address within the locale, an authorized representative | 201.7.9.1 |

201.C.5 ACCOMPANYING DOCUMENTS, instructions for use

Addition:

201.C.5.101 ACCOMPANYING DOCUMENTS, instructions for use of a CLINICAL THERMOMETER

Additional  requirements  for ACCOMPANYING  DOCUMENTS, instructions  for  use  of a CLINICAL  THERMOMETER are found in Table 201.C.103.

© ISO 2017 – All rights reserved 23

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Table 201.C.103 — ACCOMPANYING DOCUMENTS, instructions for use of a CLINICAL THERMOMETER

| Description of marking | Subclause |
| --- | --- |
| A summary of the USE SPECIFICATION | 201.7.9.2.101 a) |
| Characteristics and technical factors known to the MANUFACTURER that could pose a RISK if a single use CLINICAL THERMOMETER or its parts would be re‐used | 201.7.9.2.101 k) |
| Details of the nature and frequency of any maintenance and/or calibration needed to ensure that the CLINICAL THERMOMETER operates properly and safely | 201.7.9.2.101 i) |
| For a CLINICAL THERMOMETER that is not an ADJUSTED MODE CLINICAL THERMOMETER, the transient response | 201.101.3 |
| For a CLINICAL THERMOMETER, its parts or ACCESSORIES that contain phthalates, information on RESIDUAL RISKS for treatment of children or that of pregnant or nursing women and, if applicable, on appropriate precautionary measures | 201.11.7 |
| Information about the behaviour of the CLINICAL THERMOMETER when used without the PROBE COVER, if applicable | 201.7.9.2.101 f) 2) |
| Information concerning the disposal of the CLINICAL THERMOMETER and its components | 201.7.9.2.101 j) |
| Instructions for application of a PROBE COVER, if applicable | 201.7.9.2.101 f) 1) |
| Instructions for processing and REPROCESSING the CLINICAL THERMOMETER and its ACCESSORIES | 201.11.6.6 |
| Instructions for selection and replacement of the INTERNAL ELECTRICAL POWER SOURCE, if applicable | 201.7.9.2.101 h) |
| LABORATORY ACCURACY in the RATED OUTPUT RANGE | 201.7.9.2.101 e) |
| LABORATORY ACCURACY in the RATED EXTENDED OUTPUT RANGE, if applicable | 201.7.9.2.101 e) |
| MEASURING SITE and REFERENCE BODY SITE | 201.7.9.2.101 b) |
| RATED OUTPUT RANGE for each intended REFERENCE BODY SITE | 201.7.9.2.101 d) |
| Recommended minimum measuring time and minimum time between measurements, if applicable | 201.7.9.2.101 c) |
| RESIDUAL RISKS associated with changing environmental conditions | 201.4.2.101 |
| Whether the CLINICAL THERMOMETER is a DIRECT MODE or an ADJUSTED MODE CLINICAL THERMOMETER | 201.7.9.2.101 g) |

24 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Annex D

(informative)

Symbols on marking

IEC 60601‐1:2005+A1:2012, Annex D applies, except as follows:

Addition:

Table 201.D.2.101 — Additional symbols on marking

| No. | Symbol | Reference | Title |
| --- | --- | --- | --- |
| 1 |  | ISO 7000‐1051 Symbol 5.2.6 ISO 15223‐1:2016 | Do not re‐use |
| 2 | (YYYY‐MM) | ISO 7000‐2607 Symbol 5.1.4 ISO 15223‐1:2016 | Use by date |
| 3 |  | ISO 7000‐2499 Symbol 5.2.1 ISO 15223‐1:2016 | Sterile |
| 4 |  | ISO 7000‐2500 Symbol 5.2.2 ISO 15223‐1:2016 | Sterilized using aseptic processing techniques |
| 5 |  | ISO 7000‐2501 Symbol 5.2.3 ISO 15223‐1:2016 | Sterilized using ethylene oxide |

© ISO 2017 – All rights reserved 25

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

| No. | Symbol | Reference | Title |
| --- | --- | --- | --- |
| 6 |  | ISO 7000‐2502 Symbol 5.2.4 ISO 15223‐1:2016 | Sterilized using irradiation |
| 7 |  | ISO 7000‐2503 Symbol 5.2.5 ISO 15223‐1:2016 | Sterilized using steam or dry heat |
| 8 |  | ISO 7000‐2608 Symbol 5.2.6 ISO 15223‐1:2016 | Do not resterilize |
| 9 |  | ISO 7000‐2606 Symbol 5.2.8 ISO 15223‐1:2016 | Do not use if package is damaged |
| 10 | XXX | ISO 7000‐2725 EN 15986:2011 | Contains or presence of xxx Where PHT is used for phthalate |

26 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Annex AA

(informative)

Particular guidance and rationale

AA.1 General guidance

This annex provides a rationale for some requirements of ISO 80601‐2‐56 and is intended for those who are familiar with the subject of ISO 80601‐2‐56 but who have not participated in its development.  An understanding of the rationale underlying these requirements is considered to be essential for their  proper application. Furthermore, as clinical practice and technology change, it is believed that a  rationale will facilitate any revision of this document necessitated by those developments.

AA.2 Rationale for particular clauses and subclauses

The numbering of the following rationale corresponds to the numbering of the clauses and subclauses  in ISO 80601‐2‐56. The numbering is, therefore, not consecutive.

Subclause 201.1 Scope, object and related standards

This document addresses a range of MEDICAL ELECTRICAL EQUIPMENT (ME EQUIPMENT) generally known as  CLINICAL THERMOMETERS. This ME EQUIPMENT has been widely used to measure PATIENT temperature  during monitoring and treatment of disease. Examples of applications include detection of fever,  determination of the time of ovulation, monitoring of a physiological response to various medications  and PROCEDURES, effects of exercise and physical work, and evaluation of mental state.

CLINICAL THERMOMETERS are designed and fabricated as PORTABLE, transit‐operable, or HAND‐HELD  ME EQUIPMENT for home healthcare and clinical use, or as parts of the STATIONARY ME EQUIPMENT. The  requirements and test PROCEDURES of this document have been developed with the intent to make them  applicable to a broad range of present and future CLINICAL THERMOMETER technologies, while assuring  that every CLINICAL THERMOMETER that conforms to this document provides an acceptable degree of  diagnostic value and acceptable RISK. There are several RISKS associated with use of this type of ME EQUIPMENT. An obvious RISK is a misdiagnosis. For example, a false negative or false positive detection  of a fever which leads to a wrong treatment of a PATIENT. Another RISK is a possible injury of a PATIENT or  OPERATOR by the CLINICAL THERMOMETER or its components. RISK CONTROL is the main purpose of this document, which describes the requirements and PROCEDURES that assure acceptable levels of CLINICAL  ACCURACY and functionality, and which should be maintained over the EXPECTED SERVICE LIFE of the CLINICAL THERMOMETER.

Definition 201.3.201 ADJUSTED MODE

The  OUTPUT  TEMPERATURE  indicated  by  a  CLINICAL  THERMOMETER  is  not  necessarily  the  same  as  the  temperature  of  the  SENSOR  that  is  thermally  coupled  to  the  MEASURING  SITE.  In  the  DIRECT  MODE,  the  OUTPUT TEMPERATURE indicated by a CLINICAL THERMOMETER is the same as the temperature of the SENSOR that is thermally coupled to the MEASURING SITE. Often, DIRECT MODES are inconvenient. For example, the  time response  for an accurate measurement might be  too slow or it might be impossible  to place  the  SENSOR close to the desired body site.

For some CLINICAL  THERMOMETERS,  the OUTPUT  TEMPERATURE can be  the result of a signal adjustment or  conversion[12] and so the mode of operation is called the ADJUSTED MODE. For example, an infrared PROBE can  be  placed  in  an  ear  canal  but  the  digital  display  (OUTPUT  TEMPERATURE)  indicates  an  estimated  sublingual  temperature  of  the  PATIENT.  ADJUSTED  MODE  CLINICAL  THERMOMETERS  compensate  for

© ISO 2017 – All rights reserved 27

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

limitations  of  DIRECT  MODE  CLINICAL  THERMOMETERS  by  using  signal  processing  algorithms  to  estimate  temperature  from  measured  values,  though  there  is  often  a  corresponding  reduction  in  CLINICAL  ACCURACY. This document requires both clinical VALIDATION and laboratory VERIFICATION  for all  types of  ADJUSTED MODES.

Examples of two common types of ADJUSTED MODES follow.

Time adjustment

The  first  type  is  the  case  of  a  predictive  intermittent  CLINICAL  THERMOMETER  that  displays  an  OUTPUT  TEMPERATURE before reaching thermal equilibrium, so the displayed OUTPUT TEMPERATURE is an estimate  of  the  “would‐be”  equilibrium  temperature  by  computation,  therefore  generally  reducing  the  measurement  time.  This  type  of  ADJUSTED  MODE  is  used  to  reduce  the  measurement  time,  e.g.  from  minutes  to  several  seconds.  It  takes  about  5 min  for  a  sublingual  CLINICAL  THERMOMETER  and  about  10 min for an axillary (under the armpit) CLINICAL THERMOMETER to achieve thermal equilibrium, while a  predictive intermittent CLINICAL THERMOMETER can produce an OUTPUT TEMPERATURE in less than 20 s.

For  a  predictive  intermittent  CLINICAL  THERMOMETER,  the  rate  of  change  of  the  SENSOR'S  output  is  measured  and  used  to  predict  (anticipate)  the  final  equilibrium  temperature  which is  never  reached  during the time of measurement[13]. Figure AA.101 illustrates this type of adjustment.

Location adjustment

The second type of ADJUSTED MODE performs a positional dependent adjustment to another position that  is different from the MEASURING SITE. This position can be a physiological offset with the aim of providing  for  convenient  temperature measurement  of  a  desired  body  site  that is  different  from  the  MEASURING  SITE or can be a temperature inside the patient directly below the SENSOR.

EXAMPLE 1  A constant‐heat flow CLINICAL THERMOMETER such as a sublingual equilibrium CLINICAL THERMOMETER.

EXAMPLE 2  A radiance ear canal CLINICAL THERMOMETER that estimates the temperature at a different body site  by calculation.

EXAMPLE 3  A heat‐flow sensor at the forehead.

28 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Key

Ta  is the ambient temperature

Tm  is the MEASURING SITE temperature

TR  is the REFERENCE BODY SITE temperature

t1  is the time of thermal coupling

t2  is the time when the temperature at the MEASURING SITE is estimated

t3  is the time when achieving thermal equilibrium could have been reached

NOTE  The MEASURING SITE is the same as the REFERENCE BODY SITE.

Figure AA.101 — Example of temperature time adjustment for a predictive intermittent  CLINICAL THERMOMETER

Definition 201.3.206 CLINICAL THERMOMETER

Any  CLINICAL  THERMOMETER  (see  Figure AA.102)  contains  at  least  two  essential  components:  a  SENSOR (for example, a thermistor or a thermopile) and an output means (for example, a digital display, speaker  or  printer).  A  SENSOR  is  usually  incorporated  inside  or  positioned  near  the  PROBE.  The  PROBE  can  be  attached directly to the ENCLOSURE of a CLINICAL THERMOMETER or connected to it by a cable. To measure  temperatures,  the  PROBE is  placed  so  that  the  SENSOR is  thermally  coupled  to  the  MEASURING  SITE.  The  SENSOR  converts  thermal  energy  into  an  electrical  signal  from  which  an  output  representation  of  the  temperature is derived. Other auxiliary components and a signal processing circuit can also be included.  The auxiliary components and processing circuit can include  the ambient  temperature  SENSOR, optical  components, microcontroller, power supply and other components. A microcontroller uses a software  that  processes  the  signals  that  are  received  from  the  SENSOR  according  to  an  algorithm  that  displays  OUTPUT TEMPERATURE in the appropriate format.

© ISO 2017 – All rights reserved 29

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Key

1  MEASURING SITE

2  thermal coupling (contact or non‐contact)

3  SENSOR

4  auxiliary components

5  processing circuit

6  output means that indicates OUTPUT TEMPERATURE

Figure AA.102 — General structure of a CLINICAL THERMOMETER

Definition 201.3.207 DIRECT MODE

A DIRECT MODE CLINICAL THERMOMETER (or the TEST MODE of an ADJUSTED MODE CLINICAL THERMOMETER) is a  CLINICAL THERMOMETER whose OUTPUT TEMPERATURE is intended to represent the true temperature of the  MEASURING  SITE  or  object,  which  is  thermally  coupled  to  the  SENSOR.  In  other  words,  the  actual  temperature detected by  the SENSOR is displayed as  the OUTPUT TEMPERATURE; whatever is measured is  indicated.

Definition 201.3.219 REFERENCE BODY SITE

A CLINICAL THERMOMETER measures and indicates temperatures of a specific organ or area of a PATIENT'S body called a “body site”. The site where the temperature is actually measured (MEASURING SITE) and the  site  for  which  the  temperature  is  indicated  (REFERENCE  BODY  SITE)  are  not  necessarily  the  same.  For  example,  a  measurement  can  be  made  sublingually,  while  the  OUTPUT  TEMPERATURE  indicates  a  temperature that corresponds to the rectum. Thus, two types of “body sites” should be considered.

—  The MEASURING SITE is a place on or inside the body of the PATIENT where the PROBE is positioned and  to which the SENSOR is thermally coupled.

—  The  REFERENCE  BODY  SITE  is  the  part  of  the  body  of  the  PATIENT  whose  temperature  is  directly  measured or calculated and indicated by the output means of a CLINICAL THERMOMETER.

For  the  purpose  of  detecting  fever,  BODY  TEMPERATURES  historically  have  been  measured  by  contact  CLINICAL  THERMOMETERS in  the  sublingual,  rectal,  or  axillary  (under  the  armpit)  MEASURING  SITES which  were  the same as  the REFERENCE BODY SITES. However, most externally accessible MEASURING SITES have  not  represented  the  body  core  temperature  with  a  specific  quantitative  relationship.  Thus,  during  surgical  PROCEDURES  and  intensive  care,  temperatures  often  are  measured  by  invasive  PROBES  in  recognized  core  temperature  MEASURING  SITES  (e.g.  pulmonary  artery,  distal  oesophagus  and  urinary  bladder).  These  PROBES  are  normally  combined  with  invasive  devices,  for  example,  a  Foley  catheter.

30 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Although this document does not deny the use of new REFERENCE BODY SITES, new REFERENCE BODY SITES should  be  scientifically  and  clinically  evaluated  for  a  particular  medical  purpose.  For  ADJUSTED  MODE  CLINICAL  THERMOMETERS,  this  document  does  require  clinical  VALIDATION  with  an  existing  REFERENCE  CLINICAL  THERMOMETER.  This  requirement  means  that  the  OUTPUT  TEMPERATURE  of  a  new  CLINICAL  THERMOMETER  has  to  reference  a  REFERENCE  BODY  SITE  utilized  by  an  existing  VALIDATED  CLINICAL  THERMOMETER.

Subclause 201.4.3.101 Additional requirements for ESSENTIAL PERFORMANCE

CLINICAL  THERMOMETERS  span  the  range  from invasive  ME EQUIPMENT  with  sophisticated  ALARM  SYSTEMS that  continually  monitor  critically  ill  PATIENTS  to  simple,  inexpensive  home  healthcare  environment  ME EQUIPMENT. Every CLINICAL THERMOMETER measures or estimates the temperature of a REFERENCE BODY  SITE  for  the  purpose  of  diagnosing  or  monitoring.  These  purposes  can  be  the  detection  of  fever,  determination of the moment of ovulation, monitoring of the physiological response to medication and  PROCEDURES,  detection  of  life‐threatening  situations  (e.g.  malignant  hyperthermia,  sepsis)  and  many  other applications.

This  document  considers  it  an  unacceptable  RISK  for  a  CLINICAL  THERMOMETER  to  present  an  OUTPUT  TEMPERATURE  that is not accurate without indicating  that it is  not accurate. Methods  of indicating  this  degraded  performance  include  generating  a  TECHNICAL  ALARM  CONDITION  or  not  providing  an  OUTPUT  TEMPERATURE.  Additionally,  to  allow  for  affordable  home  healthcare  CLINICAL  THERMOMETERS,  this  document  permits  the  permissible  operating  temperature  range  to  be  marked  on  the  CLINICAL  THERMOMETER.

Subclause 201.11.6.6 Additional requirements for USABILITY of ME EQUIPMENT

The  CLINICAL  THERMOMETER,  the  PROBE  and  the  ACCESSORIES  should  not  be  operated  or  used  if  their  condition could compromise  the health or safety of  the PATIENT using  them or  the employees or  third  parties supplying them. Among other reasons, this means that the CLINICAL THERMOMETER, the PROBE and  the ACCESSORIES cannot be operated or used if there is a potential RISK of the PATIENT becoming infected  from the CLINICAL THERMOMETER, the PROBE or the ACCESSORIES.

All CLINICAL  THERMOMETERS, PROBES and ACCESSORIES are a  potential  source  of infection in  humans. Any  CLINICAL  THERMOMETERS,  PROBES  or  ACCESSORIES  that  have  been  used  by  a  PATIENT  are  potentially  contaminated  with  reproductive  human  pathogenic  microorganisms  until  proven  otherwise.  Appropriate procedures for handling and processing are essential to protect the next person handling,  or  the  next  PATIENT  using,  the  CLINICAL  THERMOMETER,  PROBE  or  ACCESSORIES.  Hence,  any  CLINICAL  THERMOMETERS,  PROBES  and  ACCESSORIES  that  have  been  used  need  processing,  following  the  MANUFACTURER’S instructions, prior to re‐use by another PATIENT.

CLINICAL THERMOMETERS, the PROBES and the ACCESSORIES if specified for re‐use and the number of re‐uses  are not limited by the specification provided by the MANUFACTURER. They will be cleaned and disinfected  several thousand times during their useful lives. When compliance with this document could be affected  by  the  cleaning  or  the  disinfecting  of  the  CLINICAL  THERMOMETER,  the  PROBE  and  the  ACCESSORIES,  an  appropriate  number  of  cycles  should  be  required  to  demonstrate  the  resistance  to  the  methods  specified  by  the  MANUFACTURER.  Considering  on  one  hand  the  test  effort  for  thousands  of  test  cycles  necessary  with  regards  to  the  expected  number  of  applications  and  on  the  other  hand  the  usually  predicted  test results if a reasonable number of cycles has been applied  to  the CLINICAL  THERMOMETER,  the  PROBE  and  the  ACCESSORIES,  the  committee  considered  a  pragmatic  number  of  100  cleaning  and  disinfection  cycles  as  appropriate  to  demonstrate  the  resistance  to  the  methods  specified  by  the  MANUFACTURER.

© ISO 2017 – All rights reserved 31

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Subclause 201.12.2.101 Additional requirements for USABILITY of ME EQUIPMENT

CLINICAL THERMOMETERS that are intended for home healthcare use (e.g. the PATIENT can be the OPERATOR)  need  to be readable by OPERATORS  that have impaired vision. Additionally, CLINICAL THERMOMETERS  that  utilize  a  segmented  display  need  an  effective  functional  test  that  ensures  that  all  segments  are  functioning. This ensures that the OPERATOR can distinguish between numbers such as 4 and 9, 5 and 6,  or 1 and 7.

Subclause 201.101.1 General test requirements

An ultimate goal of a CLINICAL THERMOMETER is to assess the true temperature of a REFERENCE BODY SITE.  Thus, CLINICAL ACCURACY of such a CLINICAL THERMOMETER can only be assessed by comparing its OUTPUT  TEMPERATURE  with  that  of  another  thermometer  (REFERENCE  THERMOMETER)  when  temperatures  are  taken concurrently  from  the same object or  the same subject. The methods of assessment are not  the  same for DIRECT MODE and ADJUSTED MODE CLINICAL THERMOMETERS. They also can be different for CLINICAL  THERMOMETERS with invasive and non‐invasive PROBES.

For  all  CLINICAL  THERMOMETERS  whose  PROBES  have  negligible  thermal  contact  with  the  environment  (zero‐heat‐flow  thermometers)  and  whose  OUTPUT  TEMPERATURES  are  not  adjusted,  a  laboratory  assessment  of performance is  sufficient. This is done by performing  temperature measurements with  the  CLINICAL  THERMOMETER  by  employing  a  REFERENCE  TEMPERATURE  SOURCE  and  comparing  the  OUTPUT  TEMPERATURE of the CLINICAL THERMOMETER with the temperature of the REFERENCE TEMPERATURE SOURCE.  This laboratory assessment is described in 201.101.

For  all  other  CLINICAL  THERMOMETERS  (ADJUSTED  MODE),  additional  clinical  VALIDATION  with  subjects  is  required,  because  the  measured  temperature  is  modified  to  indicate  the  OUTPUT  TEMPERATURE.  The  modification is a function of the physiological and anatomical properties of an average subject as well as  of the environmental conditions. Thus, to assess the effectiveness of such a CLINICAL THERMOMETER, it is  compared  with  another  CLINICAL  THERMOMETER whose  CLINICAL  ACCURACY  has  already  been  established  (i.e. a REFERENCE CLINICAL THERMOMETER). The CLINICAL ACCURACY VALIDATION is described in 201.102.

Table AA.101 summarizes the types of required testing for certain CLINICAL THERMOMETERS.

For some CLINICAL  THERMOMETERS,  the ACCOMPANYING DOCUMENT indicates  the use of a new PROBE  COVER for  every  new  measurement.  The  purpose  of  the  PROBE  COVER  is  to  minimize  the  RISK  of  cross‐ contamination between PATIENTS and to prevent soiling of the components of the CLINICAL THERMOMETER (e.g. by dust, faeces or ear wax adhering to the PROBE). Additionally, improper application of the PROBE  COVER to the CLINICAL THERMOMETER critically affects the measurement result and care should be taken to

properly  mount  or  apply  the  PROBE  COVERS  to  CLINICAL  THERMOMETERS  according  to  the  PROCEDURE indicated in the ACCOMPANYING DOCUMENT.

The disposable PROBE COVERS for some radiance (infrared) ear CLINICAL THERMOMETERS (IRETs) are made  from a polymeric material. The PROBE COVER acts as a relatively strong filter in the mid‐infrared spectral  range.  Consistent  thermal  and  optical  characteristics  of  these  disposable  PROBE  COVERS  influence  the  overall accuracy of an IRET[14]. Disposable PROBE COVERS  for CLINICAL THERMOMETERS utilizing predictive  algorithms can also affect performance. Minor  variations in  thickness  or  thermal coefficient influence  the  accuracy  of  these  CLINICAL  THERMOMETERS.  The  variability  of  the  performance  of  PROBE  COVERS influences  the  expanded  measurement  uncertainty  and  has  to  be  carefully  considered.  These  can  be  important considerations for any CLINICAL THERMOMETERS.

32 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Table AA.101 — Required tests for CLINICAL THERMOMETERS

| Type of CLINICAL THERMOMETER | LABORATORY ACCURACY VERIFICATION |  | CLINICAL ACCURACY VALIDATION |  |  |
| --- | --- | --- | --- | --- | --- |
|  | RATED OUTPUT RANGE | RATED EXTENDED OUTPUT RANGEa | CLINICAL BIAS | LIMITS OF AGREEMENT | CLINICAL REPEATABILITY |
| Intermittent or continuous CLINICAL THERMOMETER in DIRECT MODE: for example, a thermometer with invasive catheter PROBE; or a “pencil” contact thermometer for which the MEASURING SITE and REFERENCE BODY SITE are the same, e.g. sublingual, rectal, axillary (under armpit). | ✓ | ✓ | No | No | No |
| Intermittent CLINICAL THERMOMETER in ADJUSTED MODE: for example, a “pencil” predictive thermometer for which the MEASURING SITE and REFERENCE BODY SITE are not the same, e.g. sublingual, rectal, axillary (under armpit); an infrared thermometer for which the MEASURING SITE is skin or the ear canal (tympanic membrane); or a contact thermometer for which the MEASURING SITE is skin with a different REFERENCE BODY SITE. | ✓ | ✓ | ✓ | ✓ | ✓ |
| Continuous CLINICAL THERMOMETER in ADJUSTED MODE: for example, a contact thermometer for which the MEASURING SITE and REFERENCE BODY SITE are different. | ✓ | ✓ | ✓ | ✓ | No |
| a If any. |  |  |  |  |  |

Subclause 201.101.2 LABORATORY ACCURACY

The  environmental  conditions  of  ambient  temperature  and  relative  humidity  can  affect  the  performance  of  a  CLINICAL  THERMOMETER.  Within  the  RATED  range  of  such  conditions,  the  CLINICAL  THERMOMETER  is  required  to  meet  the  requirements  while  being  tested  in  a  laboratory.  Table AA.102  provides  suggested  combinations  of  REFERENCE  TEMPERATURE  SOURCE  temperature,  operating  temperature and relative humidity at which the CLINICAL THERMOMETER can be tested.

Continuous CLINICAL THERMOMETERS that are DIRECT MODE CLINICAL THERMOMETERS are intended to monitor  temperatures over time. They are well coupled to the MEASURING SITE and have a CLINICAL ACCURACY that  is equivalent  to  their  LABORATORY  ACCURACY. As a  result,  such  CLINICAL  THERMOMETERS  are  permitted  to  have a larger LABORATORY ACCURACY specification.

© ISO 2017 – All rights reserved 33

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Table AA.102 — Example combinations of operating conditions and REFERENCE temperature for  testing the LABORATORY ACCURACY

REFERENCE temperature

Upper limit of RATED OUTPUT RANGE −0,5 °C ± 0,5 °C

Midpoint of RATED OUTPUT RANGE ± 1 °C

Lower limit of RATED OUTPUT RANGE + 0,5 °C ± 0,5 °C

Operating conditions

Upper limit of RATED ambient temperature range −0,5 °C ± 0,5 °C Upper limit of RATED humidity range −5 % ± 5 %

Upper limit of RATED ambient temperature range −0,5 °C ± 0,5 °C Lower limit of RATED humidity range +5 % ± 5 %

Midpoint of RATED ambient temperature range ± 5 °C Midpoint of RATED ambient humidity range ± 20 %

Lower limit of RATED ambient temperature range +0,5 °C ± 0,5 °C Upper limit of RATED humidity range −5 % ± 5 %

Lower limit of RATED ambient temperature range +0,5 °C ± 0,5 °C Lower limit of RATED humidity range +5 % ± 5 %

The REFERENCE TEMPERATURE SOURCE is  used to check the LABORATORY ACCURACY of the CLINICAL  THERMOMETER. The largest permitted LABORATORY ACCURACY is 0,3 °C for a CLINICAL THERMOMETER that does not operate in the ADJUSTED MODE. In VERIFICATION, the expanded uncertainty of the measurement is  usually considered to be small enough if it does not exceed 1/3 of the LABORATORY ACCURACY[15]. In the  case of CLINICAL THERMOMETERS, this requires an uncertainty of the REFERENCE TEMPERATURE SOURCE at a coverage factor k = 2 of 0,3 °C/3 = 0,1 °C or less. The acceptance band has to be reduced by the  uncertainty of the reference thermometer.

Subclause 201.101.3 Time response for a CLINICAL THERMOMETER that is not an ADJUSTED MODE  CLINICAL THERMOMETER

The  ability  of  a  CLINICAL  THERMOMETER  to  track  changing  PATIENT  temperature  can  be  assessed  by  its  response  to  a  rapid  change  in  the  REFERENCE  TEMPERATURE  SOURCE.  Depending  on  the  type  of  a  continuous PROBE, the REFERENCE TEMPERATURE SOURCE can be a FLUID BATH (e.g. water bath), a BLACKBODY,  or  another  type  of  heat  source,  for  example,  a  thermally  controlled  plate.  The  test  is  performed  by  quickly  transferring  a  PROBE  from  one  REFERENCE  TEMPERATURE  SOURCE  to  another  REFERENCE  TEMPERATURE  SOURCE  that  has a different  temperature. The  faster  the OUTPUT  TEMPERATURE approaches  the REFERENCE TEMPERATURE SOURCE temperature, the better the temperature tracking ability of a CLINICAL  THERMOMETER.

Previous  standards  for  CLINICAL  THERMOMETERS  required  a  transient  response  test  for  intermittent  CLINICAL  THERMOMETERS.  These  tests  do  not  appear  in  this  document  because  ADJUSTED  MODE  CLINICAL  ACCURACY  tests  will  uncover  problems  with  response  time.  If  the  CLINICAL  THERMOMETER  is  slow  to  respond relative to the time disclosed in the instructions for use, that source of error would appear in  the resulting CLINICAL BIAS with associated LIMITS OF AGREEMENT and/or CLINICAL REPEATABILITY.

Subclause 201.102 CLINICAL ACCURACY VALIDATION

While laboratory  testing is  sufficient  for  a DIRECT  MODE  CLINICAL  THERMOMETER, it is  not  enough  for  an  ADJUSTED MODE CLINICAL THERMOMETER because adjustment algorithms are specific for the anatomical and

34 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

physiological properties of PATIENTS and the environment. These properties cannot be closely simulated  by  any  known  laboratory  REFERENCE  TEMPERATURE  SOURCE,  setting  or  instrument.  Thus,  the  CLINICAL  ACCURACY of an ADJUSTED MODE CLINICAL THERMOMETER is required to be additionally VALIDATED with actual  subjects  (PATIENTS).  Since  all  subjects  are  different,  statistical  methods  are  employed  to  compare  the  OUTPUT  TEMPERATURES  of  the CLINICAL  THERMOMETER  under  test  (DUT)  for  the REFERENCE  BODY  SITE with  those  of  a  REFERENCE  CLINICAL  THERMOMETER  (RCT)  whose  OUTPUT  TEMPERATURES  represent  the  same  REFERENCE BODY SITE as the DUT.

The RCT PROBE might not necessarily be placed at the same MEASURING SITE as the DUT PROBE. Moreover,  they both can indicate temperatures of a third site, the REFERENCE BODY SITE. For example, the DUT is a  forehead  skin  CLINICAL  THERMOMETER  (the  DUT  MEASURING  SITE  is  forehead  skin),  the  RCT  is  an  oral  CLINICAL  THERMOMETER  (the RCT  MEASURING  SITE is  sublingual)  and  they  both  output  core  temperatures  (REFERENCE  BODY  SITE  is  core).  The  RCT  CLINICAL  ACCURACY  in  representing  the  REFERENCE  BODY  SITE

temperature  was  previously  VALIDATED.  It  has  an  acceptable  value  of  CLINICAL  BIAS  that  carries  an  acceptable diagnostic RISK. If either the DUT or the RCT uses disposable PROBE COVERS, a new PROBE COVER is used  for each new measurement. This permits  the  test  to include  the PROBE  COVER  tolerances in  the  VERIFICATION of the total CLINICAL ACCURACY.

Hence, performance of a DUT in ADJUSTED MODE is assessed in two steps (see Table AA.101).  1) VERIFICATION of a LABORATORY ACCURACY as described in 201.101.

2) VALIDATION  of  CLINICAL  ACCURACY  while  taking  temperatures  from  a  sufficiently  large  group  of subjects which is described in 201.102.

Subclause 201.102.2 Human subject population requirements

There are no known data that would suggest that the CLINICAL ACCURACY of a CLINICAL THERMOMETER in an  ADJUSTED MODE is affected by the gender or race of a PATIENT. However, the age of a PATIENT is a known  and important factor[16][17]. Thermal control in humans develops as the individual grows. Besides, with  age, organs change in size, tissue texture, blood perfusion and other properties. All these can affect the  adjustment  algorithms.  When  assessing  CLINICAL  ACCURACY,  selection  of  subjects for  the  test  needs  to  include  a  sufficient  number  of  subjects  with  the  ages  for  which  the  DUT  is  intended.  This  document  recommends that four age groups be tested (see Table 201.102).

If a CLINICAL THERMOMETER is intended for a subject of any age, all four groups are required to be tested  separately. Generally, it is preferable that ages of subjects are distributed within a group as uniformly as  possible. The total number of valid subjects in any group is required to be no less than 35 and the total  number of the tested subjects is required to be at least 105. However, if a particular DUT is not intended  for all ages, the total number of tested subjects is still required to be at least 105. For example, if a basal  CLINICAL THERMOMETER is intended for ovulating females, it would be tested only with Group C. The total  number of subjects in that group is still required to be at least 105. However, there is no need to test too  young or too old female subjects. For that particular type of CLINICAL THERMOMETER, the subjects should  be females between 15 years and 50 years of age and there is no need to test older or younger subjects.

For VALIDATION  of an ADJUSTED  MODE  continuous CLINICAL  THERMOMETER,  no  separation  of age groups is  required.  However,  the  ages  of  the  subjects  should  be  uniformly  distributed  within  the  tested  population. The total number of tested subjects is still required to be at least 105.

Subclause 201.102.3 CLINICAL BIAS calculation

To  evaluate  a  CLINICAL  BIAS,  a  sufficiently large  number  of  temperature  pairs  should  be  taken  by  two  thermometers,  the  CLINICAL  THERMOMETER  under  test  (DUT)  and  the  REFERENCE  CLINICAL  THERMOMETER (RCT),  from  multiple  subjects  with  at  least  one  pair  of  OUTPUT  TEMPERATURES  per  subject.  See  Figure AA.103.

© ISO 2017 – All rights reserved 35

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Since three consecutive temperatures are taken from the MEASURING SITE by the DUT, only the first one is  used  in  computation  of  the  CLINICAL  BIAS,  because  in  clinical  practice,  typically  only  one  temperature  measurement is performed.

Key

1  line of equality

2  linear regression line

TDUT  is the OUTPUT TEMPERATURE of the DUT

TRCT  is the OUTPUT TEMPERATURE of the RCT

Figure AA.103 — Example of a comparison plot for DUT and RCT

Subclause 201.102.4 LIMITS OF AGREEMENT calculation

Dispersion of OUTPUT TEMPERATURES around the CLINICAL BIAS can be estimated by the standard deviation  σd of  the  temperature differences between  the CLINICAL THERMOMETER under  test (DUT) and  the RCT as  measured  from multiple subjects. For a graphical illustration and analysis, it is convenient  to plot  the  differences,  ΔT,  against  the  average  values  of  OUTPUT  TEMPERATURES  from  the  DUT  and  the  RCT,  i.e.  (TDUT + TRCT)/2, as shown in Figure AA.104.

If  the differences are normally distributed, it can be expected  that about 95 % of  them will  fall in  the  range  of  d  ± 2 × σd.  This  range  is  called  the  “95 %  limits  of  the  agreement”.  This  document  defines

36 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

LIMITS  OF  AGREEMENT  as  2 × σd.  It  characterizes  the  range  within  which  most  differences  between  the  OUTPUT TEMPERATURES by the DUT and the RCT lie.

Key

1  d  + 2 × σd

2  d  is the mean difference between OUTPUT TEMPERATURES from the DUT and the RCT

3  d  − 2 × σd

ΔT is the temperature differences between the CLINICAL THERMOMETER under test (DUT) and the RCT T is the average of OUTPUT TEMPERATURES from the DUT and the RCT

Figure AA.104 — Example of a Bland-Altman Plot[18] of the temperature difference (DUT minus  RCT) versus the average OUTPUT TEMPERATURES of two thermometers

Table AA.103 exemplifies  results  of a  CLINICAL  ACCURACY  test, with  the  total  number  of  PATIENTS in  the  group n = 36,  both  febrile  and  afebrile.  Three  sequential OUTPUT  TEMPERATURES  from  each  PATIENT  are  measured and indicated by a DUT as T1, T2 and T3. Note that only one DUT temperature, T1, was taken  from febrile PATIENTS (subject numbers 3 and 5). The bias and standard deviation were calculated from  the  temperature  differences  (T1 − TRCT)  as  +0,07 °C  and  0,37 °C,  respectively.  The  bias  is  acceptable.  Although the LIMITS OF AGREEMENT (0,44 °C) are on the high side, they could be considered by some to be  clinically  acceptable.  The  CLINICAL  REPEATABILITY  of  all  subjects,  other  than  febrile  subjects  less  than  5 years old, as computed using Formula (5), is σr = 0,23 °C. This number appears to be reasonable. As a  result of the test, the DUT could be considered a clinically useful thermometer.

© ISO 2017 – All rights reserved 37

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Table AA.103 — Example of CLINICAL ACCURACY VALIDATION test results

| PATIENT | T1 °C | T2 °C | T3 °C | TRCT °C | T1 − TRCT °C | σj °C |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 36,5 | 36,7 | 36,2 | 36,4 | 0,1 | 0,25 |
| 2 | 37,1 | 37,0 | 36,8 | 36,9 | 0,2 | 0,15 |
| 3 | 39,9 |  |  | 39,3 | −0,4 |  |
| 4 | 35,8 | 35,9 | 36,2 | 35,8 | 0,0 | 0,21 |
| 5 | 38,2 |  |  | 38,0 | 0,2 |  |
| 6 | 37,4 | 37,1 | 36,9 | 37,1 | 0,3 | 0,25 |
| … | … | … | … | … | … | … |
| 34 | 36,7 | 36,5 | 36,4 | 36,5 | 0,2 | 0,15 |
| 35 | 36,6 | 36,9 | 36,7 | 36,8 | −0,2 | 0,15 |
| 36 | 36,1 | 36,3 | 36,3 | 36,0 | 0,1 | 0,12 |
| Bias, d |  |  |  |  | 0,07 |  |
| Standard deviation |  |  |  |  | 0,22 |  |
| LIMITS OF AGREEMENT, LA |  |  |  |  | 0,44 |  |
| CLINICAL REPEATABILITY, σr |  |  |  |  |  | 0,23 |

Subclause 201.102.5 CLINICAL REPEATABILITY calculation

CLINICAL  REPEATABILITY  (sometimes  called  the  perceived  accuracy)  is  a  measure  of  the  consistency  of  repeated measurements with identical operating conditions when temperatures are taken within short  time intervals  from  the same MEASURING SITE of  the same subject by  the same OPERATOR with  the same  intermittent  CLINICAL  THERMOMETER.  This  characteristic  is  applicable  only  to  intermittent  CLINICAL  THERMOMETERS. To determine  the CLINICAL REPEATABILITY,  three OUTPUT  TEMPERATURES  (T1,j, T2,j, and T3,j)  are  taken  from each subject by  the CLINICAL  THERMOMETER under  test  (DUT). No RCT is required as  the  DUT  OUTPUT  TEMPERATURES  are  compared  with  themselves.  Reasonable  time  delays  between  the  measurements are necessary to minimize effects of cooling the MEASURING SITE by the PROBE and thermal  drifts  in  the  PROBE.  The  ACCOMPANYING  DOCUMENT  recommends  the  minimum  time  between  measurements  (the  highest  possible  rate  of  taking  temperatures)  at  the  MEASURING  SITE.  The  CLINICAL  REPEATABILITY  does  not  show  the  CLINICAL  ACCURACY  of  the  DUT,  only  the  consistency  of  the  OUTPUT  TEMPERATURES.

CLINICAL REPEATABILITY is a pooled standard deviation of triplicate measurements over the population of  subjects, other than febrile subjects less than 5 years old.

Subclause 201.103 PROBES, PROBE CABLE EXTENDERS and PROBE COVERS

PROBES,  PROBE  CABLE  EXTENDERS  and  PROBE  COVERS  are  as  important  in  establishing  the  safety  and  accuracy of  the complete CLINICAL  THERMOMETER as is  the indicating unit and any software algorithms.  This  subclause  establishes  that  the  MANUFACTURER  of  the  PROBE,  PROBE  CABLE  EXTENDER  or  PROBE  COVER (including a MANUFACTURER of a REPROCESSED PROBE, PROBE CABLE EXTENDER or PROBE COVER) is responsible  not  only  for  the  separately  testable  properties  (such  as  biocompatibility)  of  the  PROBE,  PROBE  CABLE  EXTENDER  or  PROBE  COVER  itself,  but  also  for  the  affected  combined  properties  (such  as  accuracy,  electromagnetic  compatibility,  and  electrical  safety)  of  the  CLINICAL  THERMOMETERS  that  the  MANUFACTURER  specifies  that  the PROBE, PROBE  CABLE  EXTENDER  or PROBE  COVER  can  be  used with. As an  example of a possible effect of REPROCESSING on biocompatibility, glutaraldehyde sterilization of silicone  rubber  materials  can  result  in  impregnation  of  the  material  with  solvent  which,  if  not  sufficiently

38 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

removed by subsequent processing, can cause a chemical burn when that PROCESS is not described (and  therefore VALIDATED) in the ACCOMPANYING DOCUMENTS[19][20][21][22].

© ISO 2017 – All rights reserved 39

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

Annex BB

(informative)

Reference temperature source

BB.1 Type of CLINICAL THERMOMETER

BB.1.1 For contact CLINICAL THERMOMETER

BB.1.1.1 REFERENCE THERMOMETER

A REFERENCE THERMOMETER, with an expanded uncertainty (coverage factor k = 2) in OUTPUT TEMPERATURE not greater than ±0,02 °C, is used to  determine the temperature of the FLUID BATH. Its  calibration  is  required to be traceable to national measurement standards.

NOTE  The definition of the coverage factor “k” can be found in ISO/IEC Guide 99[23].  BB.1.1.2 FLUID BATH

A  FLUID  BATH,  well  regulated  and  stirred  and  containing  at  least  5 l  in  volume,  is  used  to  establish  reference  temperatures  over  the  measuring  range.  It  should  be  controlled  to  have  a  suitable  temperature  stability  of  better  than  ±0,02 °C  over  the  RATED  measuring  range  of  temperature  of  the  CLINICAL THERMOMETER to be tested. It should have a temperature gradient of not greater than ±0,01 °C  within its working space at any temperature used.

BB.1.2 For non-contact CLINICAL THERMOMETER

BB.1.2.1  REFERENCE THERMOMETER and FLUID BATH described in BB.1.1.1 and BB.1.1.2.  BB.1.2.2  BLACKBODY cavity with an emissivity close to one.

References [3], [12] and [16] have examples of suitable designs of a BLACKBODY radiator.  BB.2 Uncertainty of a REFERENCE TEMPERATURE SOURCE

Total expanded uncertainty of a REFERENCE TEMPERATURE SOURCE, URTS, is calculated using Formula (B.1).

2 2 2 22

RT stab hom ε T U U U U UU Δ = + + ++ (B.1) RTS

where

URT  is an expanded uncertainty of the REFERENCE THERMOMETER;

Ustab  is an expanded uncertainty of the temperature stability of a REFERENCE TEMPERATURE SOURCE;

Uhom  is an expanded uncertainty of the temperature homogeneity (uniformity) of the REFERENCE  TEMPERATURE SOURCE;

Uε  is an expanded uncertainty of a REFERENCE TEMPERATURE SOURCE due to uncertainty of  emissivity of the BLACKBODY;

UΔT  is an expanded uncertainty due to the temperature difference between the BLACKBODY  REFERENCE THERMOMETER and the inner surface temperature of the BLACKBODY.

40 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

NOTE 1  The uncertainty components Uε and UΔT are only applicable when a BLACKBODY is utilized.  NOTE 2  In some cases, additional uncertainty components can be applicable.

In some cases, a requirement  for an expanded uncertainty of a REFERENCE  TEMPERATURE SOURCE can be  more  severe  than  0,07 °C.  An  example  is  an  accuracy  requirement  for  a  basal  CLINICAL  THERMOMETER,  which  is  0,1 °C.  In  this  case,  the  URTS  should  be  of  the  order  of  0,03 °C.  Accordingly,  the  uncertainty  components URT, Ustab and Uhom need to be smaller.

© ISO 2017 – All rights reserved 41

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Annex CC

(informative)

Reference to the essential principles of safety and performance of medical  devices in accordance with ISO 16142-1[24]

This document has been prepared to support the essential principles of safety and performance of a  CLINICAL THERMOMETER as a medical device according to ISO 16142‐1[24]. This document is intended to be  acceptable for conformity assessment purposes.

Compliance with this document provides one means of demonstrating conformance with the specific  essential principles of ISO 16142‐1[24]. Other means are possible. Table CC.1 maps the clauses and  subclauses of this document with the essential principles of ISO 16142‐1.

Table CC.1 — Correspondence between the essential principles and this document

| Essential principle of ISO 16142-1[24] | Corresponding clause(s)/subclause(s) of this document | Qualifying remarks/notes |
| --- | --- | --- |
| 1 | All | The part relating to manufacturing is not addressed. |
| a) | 206 |  |
| b) | 201.12.2, 206 |  |
| 2 | 201.4 | The part relating to manufacturing is not addressed. |
| a) | All |  |
| b) | 201.4 | The part relating to manufacturing is not addressed. |
| c) | 201.3.3, 201.7 |  |
| d) | 201.7 |  |
| 3 | All | The part relating to manufacturing is not addressed. |
| 4 | All |  |
| 5 | 201.4, 201.15 |  |
| 6 | 201.4 |  |
| 7.1 | 201.102 | Only partially covered for ADJUSTED MODE CLINICAL THERMOMETERS. The part relating to benefit‐RISK is not covered. |
| 7.2 | 201.102.1 |  |
| 8.1 | — |  |
| a) | 201.11 |  |
| b) | 201.4, 201.11 |  |
| c) | 201.9, 201.15 |  |

42 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

| Essential principle of ISO 16142-1[24] | Corresponding clause(s)/subclause(s) of this document | Qualifying remarks/notes |
| --- | --- | --- |
| 8.3 | 201.11 |  |
| 8.4 | 201.11 |  |
| 8.5 | 211, 212 |  |
| 9.1 | 201.11.6.6 |  |
| 12.1 | 201.7, 201.14, 201.16, 201.103 |  |
| 12.2 |  |  |
| a) | 206 |  |
| b) | 201.12.2, 206 |  |
| c) | 201.5, 202, 211, 212 |  |
| d) | 201.11 |  |
| e) | 201.14, 201.16 |  |
| f) | 201.11 |  |
| g) | 202 |  |
| 12.4 | 201.11 |  |
| 12.5 | 201.7, 201.8 |  |
| 12.6 | 201.1.3 |  |
| 13.1 | 201.101, 201.102, 201.103 |  |
| 13.2 | 201.7 |  |
| 13.3 | 201.12.2, 206 |  |
| 13.4 | 201.7.2.101 a), 201.7.4.3.101 |  |
| 15.1 | 201.14 |  |
| 15.2 | 201.14 |  |
| 16.1 | 201.4, 201.13 |  |
| 16.5 | 202 |  |
| 16.6 | 202 |  |
| 16.7 | 201.8 |  |
| 17.1 | 201.9 |  |
| 17.2 | 201.9 |  |
| 17.3 | 201.9 |  |
| 17.4 | 201.15 |  |
| 17.5 | 201.15 |  |
| 17.6 | 201.11 |  |
| 19.1 | 201.7 |  |
| 19.2 | 201.12.2, 206 |  |
| 20.1 | 211 | The part relating to manufacturing is not addressed. |

© ISO 2017 – All rights reserved 43

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

| Essential principle of ISO 16142-1[24] | Corresponding clause(s)/subclause(s) of this document | Qualifying remarks/notes |
| --- | --- | --- |
| 20.2 | 211 | The part relating to manufacturing is not addressed. |
| 21.1 | 201.7 |  |
| 21.2 | 201.7 |  |
| 21.3 | 201.7 |  |
| 21.4 | 201.7 |  |
| 21.5 | — |  |
| b) | 201.7.2.1.101 b) |  |
| c) | 201.7.2.1.101 c) |  |
| d) | 201.7.2.2 |  |
| e) | 201.7.2.2 |  |
| f) | 201.7.2.1, 201.7.2.1.101 |  |
| i) | 201.7.2.1.101 e) |  |
| j) | 201.7.2.1.101 e) |  |
| l) | 201.7.2.2 |  |
| m) | 201.7.2.1.101 c) |  |
| 21.6 | 201.7.2.2 |  |
| 21.7 | — |  |
| a) | 201.7.9.1 |  |
| b) | 201.7.9.1 |  |
| d) | 201.7.9.2.101 k) |  |
| k) | 201.14, 201.16 |  |
| l) | 201.7.9.2.101 i) |  |
| p) | 201.11, 211 |  |
| q) | 201.7 |  |
| 21.8 | 201.7 |  |
| 21.9 | — |  |
| a) | 211 |  |
| b) | 211 |  |
| d) | 201.1.3, 201.7.9.2.101 j) |  |
| f) | 201.7.9.2.101 e), 211 |  |

44 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Annex DD

(informative)

Terminology — Alphabetized index of defined terms

NOTE  The  ISO  Online  Browsing  Platform  (OBP)  and  the  IEC Electropedia  provide  access  to  many  of  these  terms and definitions.

| Term | Source |
| --- | --- |
| ACCESSIBLE PART | IEC 60601–1:2005+A1:2012, 3.2 |
| ACCESSORY | IEC 60601‐1:2005+A1:2012, 3.3 |
| ACCOMPANYING DOCUMENT | IEC 60601‐1:2005+A1:2012, 3.4 |
| ADJUSTED MODE | 201.3.201 |
| ALARM CONDITION | IEC 60601–1:2005+A1:2012, 3.141 |
| ALARM LIMIT | IEC 60601–1‐8:2006, 3.3 |
| ALARM SIGNAL | IEC 60601–1‐8:2005, 3.9 |
| ALARM SYSTEM | IEC 60601–1‐8:2006, 3.11 |
| APPLIED PART | IEC 60601–1:2005, 3.8 |
| BASIC SAFETY | IEC 60601–1:2005, 3.10 |
| BLACKBODY | 201.3.202 |
| BODY TEMPERATURE | 201.3.203 |
| CLEARLY LEGIBLE | IEC 60601–1:2005+A1:2012, 3.15 |
| CLINICAL ACCURACY | 201.3.204 |
| CLINICAL BIAS [Δcb] | 201.3.205 |
| CLINICAL REPEATABILITY [σr] | 201.3.206 |
| CLINICAL THERMOMETER | 201.3.207 |
| DIRECT MODE | 201.3.208 |
| ENCLOSURE | IEC 60601–1:2005, 3.26 |
| ESSENTIAL PERFORMANCE | IEC 60601–1:2005+A1:2012, 3.27 |
| EXPECTED SERVICE LIFE | IEC 60601–1:2005+A1:2012, 3.28 |
| EXTENDED OUTPUT RANGE | 201.3.209 |
| FLUID BATH | 201.3.210 |
| HAND‐HELD | IEC 60601–1:2005+A1:2012, 3.37 |
| HAZARD | IEC 60601–1:2005+A1:2012, 3.39 |
| HAZARDOUS SITUATION | IEC 60601–1:2005+A1:2012, 3.40 |
| IMMUNITY | IEC 60601–1‐2:2014, 3.8 |
| INTENDED USE | IEC 60601–1:2005+A1:2012, 3.44 |
| INTERNAL ELECTRICAL POWER SOURCE | IEC 60601–1:2005, 3.45 |

© ISO 2017 – All rights reserved 45

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

| Term | Source |
| --- | --- |
| LABORATORY ACCURACY | 201.3.211 |
| LIMITS OF AGREEMENT [LA] | 201.3.212 |
| MANUFACTURER | IEC 60601–1:2005+A1:2012, 3.55 |
| MEASURING SITE | 201.3.213 |
| ME EQUIPMENT (MEDICAL ELECTRICAL EQUIPMENT) | IEC 60601–1:2005, 3.63 |
| ME SYSTEM (MEDICAL ELECTRICAL SYSTEM) | IEC 60601–1:2005, 3.64 |
| MODEL OR TYPE REFERENCE | IEC 60601–1:2005, 3.66 |
| NORMAL CONDITION | IEC 60601–1:2005, 3.70 |
| NORMAL USE | IEC 60601–1:2005+A1:2012, 3.71 |
| OBJECTIVE EVIDENCE | IEC 60601–1:2005+A1:2012, 3.72 |
| OPERATING MODE | 201.3.214 |
| OPERATOR | IEC 60601–1:2005, 3.73 |
| OUTPUT RANGE | 201.3.215 |
| OUTPUT TEMPERATURE | 201.3.216 |
| PATIENT | IEC 60601–1:2005+A1:2012, 3.76 |
| PHYSIOLOGICAL ALARM CONDITION | IEC 60601–1‐8:2006, 3.31 |
| PORTABLE | IEC 60601–1:2005+A1:2012, 3.85 |
| PRIMARY OPERATING FUNCTION | IEC 62366–1:2015, 3.13 |
| PROBE | 201.3.217 |
| PROBE CABLE EXTENDER | 201.3.218 |
| PROBE COVER | 201.3.219 |
| PROCEDURE | IEC 60601–1:2005+A1:2012, 3.88 |
| PROCESS | IEC 60601–1:2005+A1:2012, 3.89 |
| PROGRAMMABLE ELECTRICAL MEDICAL SYSTEMS (PEMS) | IEC 60601–1:2005, 3.90 |
| RATED (value) | IEC 60601–1:2005, 3.97 |
| REFERENCE BODY SITE | 201.3.220 |
| REFERENCE CLINICAL THERMOMETER (RCT) | 201.3.221 |
| REFERENCE TEMPERATURE SOURCE | 201.3.222 |
| REFERENCE THERMOMETER | 201.3.223 |
| REPROCESSING (REPROCESSED) | 201.3.224 |
| RESIDUAL RISK | IEC 60601–1:2005+A1:2012, 3.100 |
| RESPONSIBLE ORGANIZATION | IEC 60601–1:2005, 3.101 |
| RISK | IEC 60601–1:2005+A1:2012, 3.102 |
| RISK CONTROL | IEC 60601–1:2005+A1:2012, 3.105 |
| RISK MANAGEMENT | IEC 60601–1:2005+A1:2012, 3.107 |
| RISK MANAGEMENT FILE | IEC 60601–1:2005+A1:2012, 3.108 |
| SENSOR | 201.3.225 |

46 © ISO 2017 – All rights reserved

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

| Term | Source |
| --- | --- |
| SKIN TEMPERATURE | 201.3.226 |
| SINGLE FAULT CONDITION | IEC 60601–1:2005+A1:2012, 3.116 |
| STATIONARY | IEC 60601–1:2005+A1:2012, 3.118 |
| TECHNICAL ALARM CONDITION | IEC 60601–1‐8:2006, 3.36 |
| TEST MODE | 201.3.227 |
| USABILITY | IEC 62366–1:2015, 3.16 |
| USABILITY ENGINEERING FILE | IEC 62366–1:2015, 3.18 |
| USE SPECIFICATION | IEC 62366–1:2015, 3.23 |
| VALIDATION (VALIDATED) | 201.3.228 |
| VERIFICATION (VERIFIED) | IEC 60601–1:2005+A1:2012, 3.138 |

© ISO 2017 – All rights reserved 47

BS EN ISO 80601-2-56:2017

ISO 80601-2-56:2017(E)

Bibliography

[1] IEC 60601‐1‐9, Medical electrical equipment — Part 1-9: General requirements for basic safety and essential performance — Collateral Standard: Requirements for environmentally conscious design + A1:2013

[2] HOUDAS Y. et al. Human body temperature. Plenum, New York, NY, 1982

[3] ASTM E1965‐98, Standard Specification for Infrared Thermometers for Intermittent Determination of Patient Temperature

[4] IEC 80601‐2‐591, Medical electrical equipment — Part 2-59: Particular requirements for basic safety and essential performance of screening thermographs for human febrile temperature screening

[5] IEC 60601‐1‐3, Medical electrical equipment — Part 1-3: General requirements for basic safety and essential performance — Collateral Standard: Radiation protection in diagnostic X-ray equipment

[6] IEC 60601‐2‐19:2009, Medical electrical equipment — Part 2-19: Particular requirements for the basic safety and essential performance of infant incubators

[7] ISO 9000:2015, Quality management systems — Fundamentals and vocabulary [8] ISO 14159, Safety of machinery — Hygiene requirements for the design of machinery

[9] BACHERT C., CHUCHALIN A.G., EISEBITT R., NETAYZHENKO V.Z., VOELKER M. Aspirin compared with acetaminophen in the treatment of fever and other symptoms of upper respiratory tract infection in adults: a multicenter, randomized, double‐blind, double‐dummy, placebo‐controlled, parallel‐group, single‐dose, 6‐hour dose‐ranging study. Clin. Ther. 2005, 27 (7) pp. 993–1003

[10] MULLINS M.E., EMPEY M., JARAMILLO D., SOSA S., HUMAN T., DIRINGER M.N. A prospective randomized study to evaluate the antipyretic effect of the combination of acetaminophen and ibuprofen in neurological ICU patients. 2011, 15(3), pp. 375–378

[11] Guidance for Industry and FDA Staff — Premarket Assessment of Pediatric Medical Devices. [accessed 18 July 2016]. Available from:

http://www.fda.gov/medicaldevices/deviceregulationandguidance/guidancedocuments/ucm08 9740.htm

[12] EN 12470‐5, Clinical thermometers — Part 5: Performance of infra-red ear thermometers (with maximum device)2

[13] IKEDA M. et al. U.S. Patent No. 5066141, Electronic clinical thermometer (1991‐11‐19)

[14] PUSNIK I. et al. Infrared ear thermometers — parameters influencing their reading and accuracy. Physiol. Meas. 2005, 26 pp. 1075–1084

[15] SOMMER K.D. et al. Role of measurement uncertainty in deciding conformance in legal metrology. OIML Bull. 2002, XLIII (2), pp. 19–24

1   Under preparation. Stage at time of publication: IEC DIS 80601‐2‐59:2016.

2   Withdrawn.

48 © ISO 2017 – All rights reserved

[16] JIS T4207, Infrared ear thermometers

BS EN ISO 80601-2-56:2017ISO 80601-2-56:2017(E)

[17] ERICKSON R.S. et al. Accuracy of infrared ear thermometry and other temperature methods in adults. Am. J. Crit. Care. 1994, 3 pp. 40–54

[18] BLAND J.M., ALTMAN D.G. Agreement between methods of measurement with multiple observations per individual. J. Biopharm. Stat. 2007, 17 pp. 571–582

[19] Frequently‐Asked‐Questions about the Reprocessing and Reuse of Single‐Use Devices by Third‐ Party and Hospital Reprocessors; Final Guidance for Industry and FDA Staff, July 6, 2001. [accessed 18 July 2016]. Available from:

http://www.fda.gov/medicaldevices/deviceregulationandguidance/guidancedocuments/ucm07 0864.htm

[20] Guidance for Industry and FDA Reviewers — Content and Format of Premarket Notification [510(k)] Submissions for Liquid Chemical Sterilants/High Level Disinfectants, January 3, 2000. [accessed 18 July 2016]. Available from:

http://www.fda.gov/medicaldevices/deviceregulationandguidance/guidancedocuments/ucm07 3773.htm

[21] Guidance for Industry and FDA Staff — Compliance with Section 301 of the Medical Device User Fee and Modernization Act of 2002, as amended — Prominent and Conspicuous Mark of Manufacturers on Single‐Use Devices, May 1, 2006. [accessed 18 July 2016]. Available from: http://www.fda.gov/medicaldevices/deviceregulationandguidance/guidancedocuments/ucm07 0934.htm

[22] Guidance for Industry and FDA Staff — Medical Device User Fee and Modernization Act of. 2002, Validation Data in Premarket Notification Submissions (510(k)s) for Reprocessed Single‐Use Medical Devices, September 25, 2006. [accessed 18 July 2016]. Available from: http://www.fda.gov/medicaldevices/deviceregulationandguidance/guidancedocuments/ucm07 1434.htm

[23] ISO/IEC Guide 99, International vocabulary of metrology — Basic and general concepts and associated terms (VIM)

[24] ISO 16142‐1, Medical devices — Recognized essential principles of safety and performance of medical devices — Part 1: General essential principles and additional specific essential principles for all non-IVD medical devices and guidance on the selection of standards

© ISO 2017 – All rights reserved 49

This page deliberately left blank

This page deliberately left blank

NO COPYING WITHOUT BSI PERMISSION EXCEPT AS PERMITTED BY COPYRIGHT LAW

British Standards Institution (BSI)

BSI is the national body responsible for preparing British Standards and other  standards-related publications, information and services.

BSI is incorporated by Royal Charter. British Standards and other standardization  products are published by BSI Standards Limited.

About us

We bring together business, industry, government, consumers, innovators and others to shape their combined experience and expertise into standards -based solutions.

The knowledge embodied in our standards has been carefully assembled in  a dependable format and refined through our open consultation process.  Organizations of all sizes and across all sectors choose standards to help them achieve their goals.

Information on standards

We can provide you with the knowledge that your organization needs to succeed. Find out more about British Standards by visiting our website at  bsigroup.com/standards or contacting our Customer Services team or Knowledge Centre.

Buying standards

You can buy and download PDF versions of BSI publications, including British and adopted European and international standards, through our website at  bsigroup.com/shop, where hard copies can also be purchased.

If you need international and foreign standards from other Standards Development  Organizations, hard copies can be ordered from our Customer Services team.

Copyright in BSI publications

All the content in BSI publications, including British Standards, is the property  of and copyrighted by BSI or some person or entity that owns copyright in the  information used (such as the international standardization bodies) and has  formally licensed such information to BSI for commercial publication and use. Save for the provisions below, you may not transfer, share or disseminate any  portion of the standard to any other person. You may not adapt, distribute,  commercially exploit, or publicly display the standard or any portion thereof in any  manner whatsoever without BSI’s prior written consent.

Storing and using standards

Standards purchased in soft copy format:

• A British Standard purchased in soft copy format is licensed to a sole named  user for personal or internal company use only.

• The standard may be stored on more than 1 device provided that it is accessible  by the sole named user only and that only 1 copy is accessed at any one time. • A single paper copy may be printed for personal or internal company use only. Standards purchased in hard copy format:

• A British Standard purchased in hard copy format is for personal or internal  company use only.

• It may not be further reproduced – in any format – to create an additional copy.  This includes scanning of the document.

If you need more than 1 copy of the document, or if you wish to share the  document on an internal network, you can save money by choosing a subscription  product (see ‘Subscriptions’).

Reproducing extracts

For permission to reproduce content from BSI publications contact the BSI  Copyright & Licensing team.

Subscriptions

Our range of subscription services are designed to make using standards easier for you. For further information on our subscription products go to bsigroup.com/subscriptions.

With British Standards Online (BSOL) you’ll have instant access to over 55,000  British and adopted European and international standards from your desktop. It’s available 24/7 and is refreshed daily so you’ll always be up to date.  You can keep in touch with standards developments and receive substantial  discounts on the purchase price of standards, both in single copy and subscription  format, by becoming a BSI Subscribing Member.

PLUS is an updating service exclusive to BSI Subscribing Members. You will  automatically receive the latest hard copy of your standards when they’re revised or replaced.

To find out more about becoming a BSI Subscribing Member and the benefits of membership, please visit bsigroup.com/shop.

With a Multi-User Network Licence (MUNL) you are able to host standards  publications on your intranet. Licences can cover as few or as many users as you  wish. With updates supplied as soon as they’re available, you can be sure your  documentation is current. For further information, email subscriptions@bsigroup.com.

Revisions

Our British Standards and other publications are updated by amendment or revision.  We continually improve the quality of our products and services to benefit your  business. If you find an inaccuracy or ambiguity within a British Standard or other  BSI publication please inform the Knowledge Centre.

Useful Contacts

Customer Services

Tel: +44 345 086 9001

Email (orders): orders@bsigroup.com

Email (enquiries): cservices@bsigroup.com

Subscriptions

Tel: +44 345 086 9001

Email: subscriptions@bsigroup.com

Knowledge Centre

Tel: +44 20 8996 7004

Email: knowledgecentre@bsigroup.com

Copyright & Licensing

Tel: +44 20 8996 7070

Email: copyright@bsigroup.com

BSI Group Headquarters

389 Chiswick High Road London W4 4AL UK